## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1845993 | 12/08/2023 04:30 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Needs new level sensor3/4/24-Failed "inspect fluid delivery line" check for one valve | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 04:30 PM | 04/30/2024 11:30 AM |  |  | SAH6111 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYARY026 | 451492 | Medivance Inc A Bard Medical C | 5000 |  | ER- | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1857801 | 01/11/2024 09:12 AM | CM | Corrective Maintenance | IP | In Progress | RPT project OR | Crystal Schmitt | CRSCH | EMPLOYEE | 01/16/2024 09:12 AM | 03/27/2024 10:50 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1871138 | 02/15/2024 04:29 PM | CM | Corrective Maintenance | IP | In Progress | PB980 training | Crystal Schmitt | CRSCH | EMPLOYEE | 02/20/2024 04:29 PM | 02/15/2024 05:25 PM |  |  | TRAINING | Used for Training Work Orders | CC-056 | Environment | NSN | 000 | Misc-Numbers | Training |  |  | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1872943 | 02/20/2024 05:34 PM | CM | Corrective Maintenance | WP | Awaiting Parts | This device and other system components are plugged into a power transformer, which is plugged into a non-approved power strip. CM to correct power strip issue. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 05:34 PM | 02/20/2024 05:38 PM |  |  | SAH4002 | Stimulus Generator Unit | 17-898 | Monitoring Systems, Physiologic, Cardiac Electrophysiology | 4417-SGU/SC397 | 427969 | Bard Peripheral Vascular Div C | MP3008 |  | EP LAB | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1874209 | 02/26/2024 11:12 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118, Not passing negative pressure and tempLOC: Water Room, Dialysis | Crystal Schmitt | CRSCH | EMPLOYEE | 03/02/2024 11:12 AM | 02/27/2024 01:36 PM |  |  | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water Room, Dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1882410 | 03/12/2024 10:55 AM | CM | Corrective Maintenance | IP | In Progress | Need key tether installed | Crystal Schmitt | CRSCH | EMPLOYEE | 03/17/2024 10:55 AM | 03/12/2024 10:59 AM |  |  | SAH5531 | Endo Motorized Travel Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 2000RFWDII082217A | 453029 | PHS West Inc | Ergo Express |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1883634 | 03/19/2024 08:03 PM | CM | Corrective Maintenance | IP | In Progress | Found power cord/charger damaged during department PMs | Crystal Schmitt | CRSCH | EMPLOYEE | 03/24/2024 08:03 PM | 03/19/2024 08:05 PM |  |  | SAH3370 | Printer | 22-203 | Printers, Computer/Network, Ink-Jet | U64418-G6W123430 | 473839 | Brother International | MW-260MFi |  | PIC TEAM | 1105-37200 | SAH-IV Therapy | 04 | Inventory | SAH | SAH
1885219 | 03/21/2024 12:44 PM | CM | Corrective Maintenance | IP | In Progress | PM failed-Failed pulse width, pulse frequency and stimulus duration test. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/26/2024 12:44 PM | 04/18/2024 07:50 AM |  |  | SAH5393 | Electroshock Unit | 11-484 | Stimulators, Electrical, Brain, Convulsive Therapy | 42685 | 104776 | Somatics LLC | System IV |  | SAH-PACU | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH
1885387 | 03/22/2024 11:01 AM | CM | Corrective Maintenance | AD | Awaiting Delivery of Equipment | 727-825-1125 or 727-692-9317/Not running blood pressure. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 11:01 AM | 04/09/2024 04:45 PM |  |  | SAH5881 | Front End Hemodynamic Monitor | 12-648 | Monitoring Systems, Physiologic, Cardiac Catheterization | US1550001478 | 452291 | Philips Interventional Patient | FC2020 |  | Cath Lab 2 | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1885408 | 03/22/2024 12:25 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Found during PM-Bottom corner LCD is dark and corner of front chassis damaged. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 12:25 PM | 03/22/2024 12:25 PM |  |  | SAH5687 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311035HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 09 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1885424 | 03/22/2024 01:17 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to top and bottom housing | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 01:17 PM | 03/22/2024 01:01 PM |  |  | L42214 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 82323 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1885828 | 03/25/2024 03:56 PM | CM | Corrective Maintenance | IP | In Progress | 7278251287CO2 cylinder tank holder install needed, on siteEndo storage room 2 | Crystal Schmitt | CRSCH | EMPLOYEE | 03/30/2024 03:56 PM | 04/09/2024 02:30 PM |  |  | SAH5531 | Endo Motorized Travel Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 2000RFWDII082217A | 453029 | PHS West Inc | Ergo Express |  | na | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1886842 | 03/26/2024 11:54 AM | CM | Corrective Maintenance | IP | In Progress | Time did not update on large wall display | Crystal Schmitt | CRSCH | EMPLOYEE | 03/31/2024 11:54 AM | 04/15/2024 01:00 PM |  |  | SAH9311 | OR Video Integration Systems | 20-388 | Automation Systems, Operating Room | R221790457 | 107114 | STERIS Corp | HexaVue System |  | Surgery Core#3 near OR15 | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1888539 | 03/29/2024 01:44 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727 825 1118RO Machine: Service Required. Component failure | Crystal Schmitt | CRSCH | EMPLOYEE | 04/03/2024 01:44 PM | 04/30/2024 03:25 PM |  |  | SAH8862 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1407678 | 199732 | Mar Cor Purification A Cantel | MILLENIUM HX |  | SAH / Dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1889662 | 04/01/2024 10:20 AM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Housing damaged | Crystal Schmitt | CRSCH | EMPLOYEE | 04/06/2024 10:20 AM | 04/16/2024 11:15 AM |  |  | SAH9589 | Laparoscopic Insufflator | 16-849 | Insufflators, Laparoscopic | C03602 | 457057 | Medivators Corp A Cantel Medic | EGA-501 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1891643 | 04/02/2024 02:03 PM | CM | Corrective Maintenance | IP | In Progress | 7278251287- SAH Dictation room - Device giving error saying "DR3 battery needs to be replaced" | Crystal Schmitt | CRSCH | EMPLOYEE | 04/07/2024 02:03 PM | 04/09/2024 02:15 PM |  |  | SAH12058 | Data Recorder Kit | 34-626 | Recorders, Physiologic, pH, Upper Gastrointestinal Tract, Ingestible Capsule | DR30823485 | 391192 | Medtronic Given Imaging Inc | FGS-0347 |  | SAH-Endo/Dictation Room | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1892294 | 04/04/2024 06:26 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM Failed-USB port damaged-unable to block out port or connect any USB. Bottom corner of display is dark-display needs to be replaced. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/09/2024 06:26 PM | 04/09/2024 09:10 AM |  |  | SAH8034 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417321139HA | 440964 | GE Healthcare Technologies | B650 |  | EPACU01 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1892711 | 04/08/2024 11:42 AM | CM | Corrective Maintenance | IP | In Progress | Found during PM-Damaged front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 04/13/2024 11:42 AM | 04/09/2024 07:55 AM |  |  | SAH5608 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311028HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 16 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1892850 | 04/09/2024 08:39 AM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Front chassis is damaged-clip attachment is broken off. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/14/2024 08:39 AM | 04/09/2024 08:20 AM |  |  | SAH4953 | Bispectral Index BIS Module | 20-783 | Physiologic Monitor Modules, Electroencephalography | SGS19130002HA | 440964 | GE Healthcare Technologies | E-BIS-01 | CVICU | cvICU | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1893772 | 04/12/2024 06:01 AM | CM | Corrective Maintenance | AC | Acknowledged | 727-825-1118 \ Turned off with ''no water'' error \ Loc: dialysis water room | Crystal Schmitt | CRSCH | EMPLOYEE | 04/17/2024 06:01 AM | 04/15/2024 09:45 AM |  |  | SAH1865 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 9T0S236481 | 312187 | Fresenius Medical Care North A | 2008 T |  | T7 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1894348 | 04/16/2024 07:10 AM | CM | Corrective Maintenance | OF | Out For Repair | Problem: piece broken off of back of deviceLOC: EndoPhone:727.825.1287 | Crystal Schmitt | CRSCH | EMPLOYEE | 04/21/2024 07:10 AM | 04/22/2024 08:35 AM |  |  | SAH4453 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A973741 | 107875 | Exergen Corp | TAT-5000 | ENDO | Endo 2 | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1894486 | 04/16/2024 10:58 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Device has a cracked rear bezel and an issue with the display backlight. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/21/2024 10:58 AM | 04/26/2024 09:55 AM |  |  | SAH7374 | Non-Invasive Ventilator | 20-746 | Ventilators, Noninvasive Positive Pressure | 100040164 | 454161 | Philips Respironics | V60 |  | AH_V60_06 | 1105-35000 | SAH-Respiratory Care Services 2nd | LS | High Risk | SAH | SAH
1894517 | 04/16/2024 12:31 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Pressure gague broken, weak suction | Crystal Schmitt | CRSCH | EMPLOYEE | 04/21/2024 12:31 PM | 04/16/2024 12:10 PM |  |  | SAH5686 | Portable Suction | 27-657 | Aspirators, Airway | PD505288 | 453450 | DeVilbiss Healthcare Inc | 7305P-D |  | ICU in front of room 13 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1897585 | 04/19/2024 02:09 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Endo leak tester has a leak in the tubing | Crystal Schmitt | CRSCH | EMPLOYEE | 04/24/2024 02:09 PM | 04/22/2024 07:58 AM |  |  | SAH5701 | Endoscope- Pneumatic Tester | 23-981 | Testers, Endoscope, Leak/Fluid Detection | 7659553 | 276750 | Olympus America Inc | MU-1 | ENDO | Endo 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1900809 | 04/25/2024 04:25 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Power cord ground loose and ES testing failed. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/30/2024 04:25 PM | 04/25/2024 03:50 PM |  |  | SAH6334 | Surgical Imaging Table | 16-600 | Tables, Imaging, Radiographic/Fluoroscopic | ISR3065 | 170423 | Image Diagnostics Inc | ISR G2 |  | COV Bed 3 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1901324 | 04/29/2024 01:01 PM | CM | Corrective Maintenance | IP | In Progress | 727-409-4685 / Treadmill belt keeps stopping randomly, interrupting tests performed using the treadmill. Provider does not see any specific causes for the sudden stops. Treadmill does have a service sticker dated for 7-23. / 3rd Floor, Echo Stress dept. / | Crystal Schmitt | CRSCH | EMPLOYEE | 05/01/2024 01:01 PM | 04/30/2024 06:45 PM |  |  | SAH11281 | Treadmill | 23-827 | Stress Test System, Treadmill | GEDC-5244 | 440964 | GE Healthcare Technologies | T2100-ST2 |  | Echo stress dept | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH