## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | REQUESTOR | CONTACT | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | PRIORITY | PRIORITY_DESC | TAG_NUMBER | DESCRIPTION | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | FACILITY | FACILITY_DESC
1568048 | 02/09/2022 10:59 AM | SP | Special Projects | IW | Awaiting Onsite Vendor Maint | b119865 | Robert Nobrega / 7278594913 | Log4j Vulnerability Coordinate with vendor for remediation | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS6406 | Digital Portable X-ray System | DF2402000764WK | 439946 | GE Healthcare USA | 5555000-27 | SJS | SJS
1748847 | 04/19/2022 03:08 PM | SP | Special Projects | IW | Awaiting Onsite Vendor Maint | b119865 | ROBERT NOBREGA / 7278594913 | Remediate Axeda cybersecurity vulnerability following attached procedure | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC120108 | Computer Aided Detection Mammo | C5740 | 106487 | Hologic Inc | ASY-02354 | BOI | BOI
1772836 | 08/01/2023 07:46 AM | PM | Planned Maintenance | WP | Awaiting Parts |  |  | -Vendor PM - Next Scheduled Date: 02/01/2024 | KenMed Supply LLC | 569382 | VENDOR | 2 | PM Level 2 | WHH3279 | Surgical Table | 3500-3I-1286 | 102186 | Skytron Div The KMW Group | 3500 | WHH | WHH
1772907 | 08/01/2023 07:46 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint |  |  | -Vendor PM - Next Scheduled Date: 02/01/2024 | KenMed Supply LLC | 569382 | VENDOR | 2 | PM Level 2 | WHH3620 | Surgical Table | 3500-4G-1385 | 102186 | Skytron Div The KMW Group | 3500 | WHH | WHH
1772915 | 08/01/2023 07:46 AM | PM | Planned Maintenance | WP | Awaiting Parts |  |  | -Vendor PM - Next Scheduled Date: 02/01/2024 | KenMed Supply LLC | 569382 | VENDOR | 2 | PM Level 2 | WHH3629 | Surgical Table | 00259 | 456817 | MAQUET Medical Systems USA | 1425.0180 | WHH | WHH
1773865 | 06/08/2023 02:27 AM | EN | Easy Net | WP | Awaiting Parts | b144654 | 8632971805 | not connecting to program. | Med Stat Consulting SVCS | M10025 | VENDOR | 9 | 2 Routine | WHH7058 | Pump Control Unit 5.7 Screen | 15844815 | 378348 | BD | 8015LS | WHH | WHH
1777067 | 06/19/2023 09:59 AM | EN | Easy Net | IP | In Progress | b113787 | 863-297-1807 | IV pump wont scan | Med Stat Consulting SVCS | M10025 | VENDOR | 9 | 2 Routine | WHH2475 | Large Volume Pump Module | 4123275 | 378348 | BD | 8100 | WHH | WHH
1783113 | 07/06/2023 02:14 PM | CM | Corrective Maintenance | WP | Awaiting Parts | B105822 | Dianne Gee / 813-443-7163 | 8133028045 - Monitor on top broken off LOC: OR cove hallway | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4274 | Mobile C-Arm | E2-4440 | 453503 | GE Healthcare Surgery | 9900 ELITE | SJS | SJS
1805124 | 09/01/2023 08:15 AM | PM | Planned Maintenance | NP | PM Not Performed Missing |  |  | --- Ranger 245 PM Procedure --- - Next Scheduled Date: 12/01/2023 | KenMed Supply LLC | 569382 | VENDOR | 2 | PM Level 2 | WHH8867 | Blood Warmer | 3425 | 418683 | Arizant Healthcare Inc | 245 | WHH | WHH
1854698 | 01/03/2024 03:51 AM | SS | Support Services | IW | Awaiting Onsite Vendor Maint | b092528 | John Sigountos / 727-315-5053 | This is a reminder that there is a remote update required for your medical device. Please reboot and accept update on impacted functional location 400-431939 , syngo X-WP at ST. JOSEPH HOSPITAL-SOUTH , RIVERVIEW , FL .Email attached | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS6417 | Leonardo Workstation | 12162 | 399199 | Siemens Medical Solutions USA | 10281061 | SJS | SJS
1869137 | 02/06/2024 08:50 AM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | b115690 | 813-302-8743 | Siemen"s Leonardo PC will not transmit images to ISCV or Artis. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS6414 | AX Application Software SYS | 25142 | 399199 | Siemens Medical Solutions USA | 10848815 | SJS | SJS
1870088 | 02/09/2024 03:41 PM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | b115690 | 8133028743 | Grounding wire to cath lab table found disconnected. Wire is at Charge desk for cath lab. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4415 | Angiographic/Interventional Sy | 121056 | 399199 | Siemens Medical Solutions USA | 10848282 | SJS | SJS
1880455 | 03/06/2024 11:54 AM | CM | Corrective Maintenance | WP | Awaiting Parts | B117911 | Steven Kelly / | Exposure switch cable is stretched out.- New ED area (Portable #4) - 813-302-8045 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS7168 | Digital Portable X-ray System | DF2401900182WK | 439946 | GE Healthcare USA | 5555000-27 | SJS | SJS
1882305 | 03/11/2024 05:02 PM | CM | Corrective Maintenance | WP | Awaiting Parts | b117911 | Steven Kelly / | While inspecting the unit, I noticed that the ethernet port was missing. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS6406 | Digital Portable X-ray System | DF2402000764WK | 439946 | GE Healthcare USA | 5555000-27 | SJS | SJS
1883300 | 03/18/2024 09:07 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B158750 | Gerson Colon jr / 813-210-6778 | Lisa Stewart emailed me stating: "CT1 is down since Yesterday, March 17th". | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS2557 | CT Scanner 128 | 95013 | 399199 | Siemens Medical Solutions USA | 8098027 | SJS | SJS
1885030 | 03/20/2024 11:32 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B117911 | Steven Kelly / | 813.302.8045 / Problem: 14x17 Image receptor has been failing QA for a couple months / LOC: 1st Floor - X-Ray | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS7168 | Digital Portable X-ray System | DF2401900182WK | 439946 | GE Healthcare USA | 5555000-27 | SJS | SJS
1885055 | 03/20/2024 01:45 PM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | EXD53563 | 813-356-7113 | field detector out of field/ center. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH28126 | Radiographic/Fluoroscopic Syst | 41C351685001 | 106973 | Shimadzu Medical Systems, Inc | SonialVision G4 | SJH | SJH
1886801 | 03/26/2024 10:51 AM | IN | Initial Inspection | DO | Device Operational | b126836 | Jennah Judd / 727-315-5339 | REC 4/1 AFFIRM PRONE 5100-33510 BREAST CENTER 261388820 HOLOGIC QUOTE ATTACHED | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | NEW-EQUIP | New Equipment ADD | NSNNEW-EQUIPSJH | 000 | Misc-Numbers | New Equipment ADD | SJH | SJH
1886919 | 03/26/2024 03:23 PM | CM | Corrective Maintenance | WP | Awaiting Parts | B158750 | Gerson Colon jr / 813-210-6778 | Received email stating, the Barco Monitor's video card is bad and needs replacement. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4871 | 3MP Color Med Display | 2590235996 | 392478 | Barco Inc | MDNC-3421 | SJS | SJS
1891633 | 04/02/2024 01:46 PM | EN | Easy Net | IP | In Progress | b149445 | 8632971800 | Pump says channel error. Cant disconnect from pump | Med Stat Consulting SVCS | M10025 | VENDOR | 9 | 2 Routine | WHH2116 | Large Volume Pump Module | 4124694 | 378348 | BD | 8100 | WHH | WHH
1894731 | 04/17/2024 07:03 AM | CM | Corrective Maintenance | IP | In Progress | ne136866 | Mathieu Clerveaux | 813-302-8743/C-Arm touchscreen is not powering on. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS5420 | Front End Hemodynamic Monitor | US1638002577 | 452291 | Philips Interventional Patient | FC2010 | SJS | SJS
1894758 | 04/17/2024 08:38 AM | CM | Corrective Maintenance | WP | Awaiting Parts | B117911 | Steven Kelly / | 813-302-8045/Will not move. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4274 | Mobile C-Arm | E2-4440 | 453503 | GE Healthcare Surgery | 9900 ELITE | SJS | SJS
1897709 | 04/21/2024 04:57 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B133890 | Julie Dang / | 8133028045Detector not holding charge, not pairing to binImaging outside catscan | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS5488 | Digital Portable X-ray System | 1034329WK1 | 439946 | GE Healthcare USA | 5555000-5 | SJS | SJS
1898825 | 05/01/2024 06:31 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Janet Pindral | 813.872.3949 | -Service Agreement PM - Next Scheduled Date: 11/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH4067 | 2D Digital Mammo | 81008110911 | 106487 | Hologic Inc | SDM-00001-2D | SJH | SJH
1899228 | 04/22/2024 07:39 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B157579 | Seth Pini / | 813-302-8743 / Repair request for lead shield mounted to ceiling / IR 1 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4415 | Angiographic/Interventional Sy | 121056 | 399199 | Siemens Medical Solutions USA | 10848282 | SJS | SJS
1899413 | 04/22/2024 10:28 PM | CM | Corrective Maintenance | CF | Conditionally Functional | JRB57478 | Jason Bowers / 813-870-4605 | Radiology 2nd floor : 813-870-4605: Plug is missing the ground | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH26543 | Digital Portable X-ray System | DF2401900172WK | 439946 | GE Healthcare USA | 5555000-47 | SJH | SJH
1901017 | 04/27/2024 01:13 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B096367 | Maria Gonzalez / 813-870-4605 | 813-870-4000 x804605 OR x735494 -losing connection with board when placed behind patient LOC: Surgery dept 2nd floor equipment storage | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH16157 | Digital Portable X-ray System | DF2401800163WK | 439946 | GE Healthcare USA | 5555000-27 | SJH | SJH
1901050 | 04/28/2024 01:41 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B077072 | Phil Lake / 813-870-4625 | 8138704625 / Need PO order for vendor repair / Catscan | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH8157 | CT Scanner | 8DA15X2004 | 101894 | Toshiba America Medical System | TSX-301C-6 | SJH | SJH
1901370 | 04/29/2024 03:25 PM | IN | Initial Inspection | AD | Awaiting Delivery of Equipment | b126836 | Jennah Judd / 727-315-5339 | ETA 5/7/2024 3D DIMENSION 5101-33510 BREAST CENTER 510122972 HOLOGIC QUOTE ATTACHED | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | NEW-EQUIP | New Equipment ADD | NSNNEW-EQUIPSJS | 000 | Misc-Numbers | New Equipment ADD | SJS | SJS
1903490 | 05/01/2024 11:10 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Val Relente | 813-302-8030 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 05/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJS4274 | Mobile C-Arm | E2-4440 | 453503 | GE Healthcare Surgery | 9900 ELITE | SJS | SJS
1904263 | 05/01/2024 12:06 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | n | n | -Service Agreement PM - Next Scheduled Date: 11/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC186779 | MRI 1.5 Tesla | 31351 | 399199 | Siemens Medical Solutions USA | 10018165 | BOI | BOI
1904491 | 05/02/2024 09:36 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B127848 | Amanda Logan / | 813.586.8508 / Mammo machine buttons need to be replaced. The machine is still usable in the meantime. / Bloomingdale HealthHub | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC146753 | 3D Digital Mammo | 3DM160100701 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1904634 | 04/29/2024 03:25 PM | IN | Initial Inspection | IW | Awaiting Onsite Vendor Maint | b126836 | Jennah Judd / 727-315-5339 | ETA 6/6/2024 FORESIGHT TIMS COMPUTER 5102-21221 SJS-ENDOSCOPY | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | NEW-EQUIP | New Equipment ADD | NSNNEW-EQUIPSJS | 000 | Misc-Numbers | New Equipment ADD | SJS | SJS
1906023 | 05/13/2024 08:34 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B119731 | Dannielle Alfonso / | 813-259-1907/Gantry will not power on. Siemens ticket# 400112602478 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC171258 | CT Scanner | 64559 | 399199 | Siemens Medical Solutions USA | SomatomDefinition AS | BOI | BOI
1906426 | 05/14/2024 08:08 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B162030 | Kathryn Gullo / 334.494.4072 | X-Ray table's Lead shield is unravelingEquipment needs tag | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4414 | Angiographic/Interventional Sy | 109463 | 399199 | Siemens Medical Solutions USA | 10848281 | SJS | SJS
1906642 | 05/15/2024 07:15 AM | CM | Corrective Maintenance | DU | Device Unavailable | B165360 | Victoria Perez / 813-507-3032 | 813-302-8045. Cassette is not connecting. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS7168 | Digital Portable X-ray System | DF2401900182WK | 439946 | GE Healthcare USA | 5555000-27 | SJS | SJS
1906745 | 05/15/2024 12:53 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B158750 | Gerson Colon jr / 813-210-6778 | Per Jamie Larussa: "Several issues with the TIMS computer." | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH24247 | Dicom Conversion System | 7001137 | 471654 | TIMS Medical Div Foresight Ima | 061000-2000 | SJH | SJH
1906747 | 05/14/2024 01:01 PM | CM | Corrective Maintenance | WP | Awaiting Parts | B158750 | Gerson Colon jr / 813-210-6778 | Nikki Moore stated: "Ceiling monitors keep moving around." | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS2556 | CT Scanner 128 | 95021 | 399199 | Siemens Medical Solutions USA | 8098027 | SJS | SJS
1906831 | 05/16/2024 03:55 AM | CM | Corrective Maintenance | AF | Awaiting FSR | B090105 | Justin Olson / 727-825-1040 | ASAP - Not turning on after CT machine was reset. 813-302-8851 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS2556 | CT Scanner 128 | 95021 | 399199 | Siemens Medical Solutions USA | 8098027 | SJS | SJS
1911770 | 06/01/2024 10:55 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Janet Pindral | 813.872.3949 | -Service Agreement PM - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH4072 | 2D Digital Mammo | 81005110811 | 106487 | Hologic Inc | SDM-00001-2D | SJH | SJH
1912172 | 05/22/2024 10:10 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B132784 | Deyanira Rizzo / | 8133567112- -Tag#sjh22327 says invalid- TM reports compression test failed | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | NOTAG | Need Tag Number | 16B01337 | 000 | Misc-Numbers | No Tag | SJH | SJH
1912278 | 05/22/2024 04:13 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B110516 | Abby Hastings / 813-586-8511 | 813-586-8511 / Top part has come off, can''t get it back on. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC176686 | 3D Digital Mammo | 3DM160100543 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1913007 | 05/28/2024 07:48 AM | CM | Corrective Maintenance | IP | In Progress | B119731 | Dannielle Alfonso / | 813-259-1907. CT Machine is down. Reference ticket from Siemens 40011262701. | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | BC171258 | CT Scanner | 64559 | 399199 | Siemens Medical Solutions USA | SomatomDefinition AS | BOI | BOI
1913179 | 05/28/2024 01:45 PM | CM | Corrective Maintenance | IP | In Progress | B105822 | Dianne Gee / 813-443-7163 | Will not go down on column, unable to activate drive. Location: Radiology by nuke med dept. Contact: 8133028045 TM called to get update and advised they need device-8133028045 TM called and advised device is working and WO can be closed. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS6405 | Digital Portable X-ray System | 1034415WK8 | 439946 | GE Healthcare USA | 5555000-5 | SJS | SJS
1913191 | 05/28/2024 02:23 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B110516 | Abby Hastings / 813-586-8511 | 813-493-1176. Cad marking are showing up on all images. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC176686 | 3D Digital Mammo | 3DM160100543 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1913192 | 05/28/2024 02:24 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B110516 | Abby Hastings / 813-586-8511 | 813-493-1176. Cad marking on all images. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC146753 | 3D Digital Mammo | 3DM160100701 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1913483 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | Uninterruptible Power Systems - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC110005 | Uninterruptible Power Systems | EG161CBC02 | 450901 | Eaton Corp | 9390 | BOI | BOI
1913484 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | -Pm Device Using Mfg Procedure - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC110006 | Laser Imager | 09927014 | 441054 | Konica Minolta Medical Imaging | 873 | BOI | BOI
1913485 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | Uninterruptible Power Systems - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC110007 | Uninterruptible Power Systems | EG161CBC07 | 450901 | Eaton Corp | 9390 | BOI | BOI
1913487 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | -Service Agreement PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC110016 | General Digital Rad Room | 22714 | 399199 | Siemens Medical Solutions USA | 10281013 | BOI | BOI
1913490 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | Densitometer - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC119251 | Densitometer B&W Film QC | 112745 | 102741 | X-Rite Inc | 331 | BOI | BOI
1913491 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | Meredith Alaynick | TBD | -Service Agreement PM - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC120211 | Bone Densitometer | 301988M | 106487 | Hologic Inc | Horizon-W | BOI | BOI
1913499 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | Synda McGraph | TBD | -Service Agreement PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC146753 | 3D Digital Mammo | 3DM160100701 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1913528 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186802 | 3MP Color Med Display | 2590235115 | 392478 | Barco Inc | MDNC-3421 | BOI | BOI
1913530 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186803 | 3MP Color Med Display | 2590235112 | 392478 | Barco Inc | MDNC-3421 | BOI | BOI
1913531 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186804 | 3MP Color Med Display | 2590134129 | 392478 | Barco Inc | MDNC-3421 | BOI | BOI
1913532 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186806 | 3MP Color Med Display | 2590134130 | 392478 | Barco Inc | MDNC-3421 | BOI | BOI
1913533 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186807 | 12MP diagnostic display | 2590232628 | 392478 | Barco Inc | MDMC-12133 | BOI | BOI
1913534 | 06/01/2024 11:38 AM | PM | Planned Maintenance | OP | Open | n | n | -Service Agreement PM and Required Annual Password Change - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186808 | Bone Densitometer | 301016M | 106487 | Hologic Inc | Horizon-W | BOI | BOI
1913580 | 05/30/2024 11:55 AM | CM | Corrective Maintenance | IP | In Progress | B105822 | Dianne Gee / 813-443-7163 | 813-302-8045/Images not sending to PACS. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS5900 | Dicom conversion system | 8002143 | 471654 | TIMS Medical Div Foresight Ima | TIMS 2000 SP | SJS | SJS