## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1809239 | 09/18/2023 08:04 AM | CM | Corrective Maintenance | IP | In Progress | 727 502 4411 US machine, Sonisite - Error Code: P...... | Crystal Schmitt | CRSCH | EMPLOYEE | 09/23/2023 08:04 AM | 11/13/2023 09:00 AM |  |  | SAH8200 | Portable Ultrasound System | 18-143 | Scanning Systems, Ultrasonic, Portable | WK3D03 | 340440 | SonoSite div of FUJIFILM | P08778 |  | CATH LAB | 1105-30610 | SAH-Cardiac Cath Lab | 04 | Inventory | SAH | SAH
1822860 | 10/20/2023 07:44 AM | CM | Corrective Maintenance | OF | Out For Repair | 727 825 1021Broken transport monitor - fell off a shelf and is cracked@ front desk | Crystal Schmitt | CRSCH | EMPLOYEE | 10/25/2023 07:44 AM | 11/22/2023 12:40 PM |  |  | SAH12088 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT21242066SA | 440964 | GE Healthcare Technologies | B450 |  | TRANSPORT | 1105-20100 | SAH-Anesthesiology | IU | Active | SAH | SAH
1823628 | 10/24/2023 01:50 PM | CM | Corrective Maintenance | OF | Out For Repair | Reading incorrect ground resistance | Crystal Schmitt | CRSCH | EMPLOYEE | 10/29/2023 01:50 PM | 10/26/2023 09:20 AM |  |  | SAH8877 | Electrical Safety Analyzer | CC-040 | Test Equipment | 5942040 | 178848 | Fluke Biomedical | ESA609 |  |  | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1828936 | 11/01/2023 01:10 PM | CM | Corrective Maintenance | IP | In Progress | Special Projects and assignments November 2023 | Crystal Schmitt | CRSCH | EMPLOYEE | 11/06/2023 12:10 PM | 11/27/2023 05:22 PM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1829136 | 11/02/2023 12:10 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed due to physical damage, needs new front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/07/2023 11:10 AM | 11/15/2023 12:20 PM |  |  | SAH3229 | LP20e Defibrillator/Pacer | 17-882B | Defibrillator/Pacemakers, External, Biphasic | 48234721 | 452575 | Physio-Control Inc | LP20E |  | SAH-3rd EP lab | 1105-13100 | SAH-Intensive Care Unit | LS | High Risk | SAH | SAH
1830011 | 11/07/2023 06:01 AM | CM | Corrective Maintenance | AF | Awaiting FSR | 727-825-1118Loc: DialysisUnit has Air Alarm, unbale to clear | Crystal Schmitt | CRSCH | EMPLOYEE | 11/12/2023 06:01 AM | 11/27/2023 09:15 AM |  |  | SAH4193 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248892 | 312187 | Fresenius Medical Care North A | 2008 T |  | T12 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1830745 | 11/10/2023 08:17 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 7278251287 - Unit not seeing 2 devices - 2nd floor 2 SE endo 5 | Crystal Schmitt | CRSCH | EMPLOYEE | 11/15/2023 08:17 AM | 11/10/2023 01:20 PM |  |  | SAH5343 | Intergration System | 20-388 | Automation Systems, Operating Room | RL0010F367C4EA | 107114 | STERIS Corp | IQ2800 |  | Under Desk Endo 5 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1835395 | 11/13/2023 10:22 AM | CM | Corrective Maintenance | OF | Out For Repair | Device says "Reconnect PDM" error message | Crystal Schmitt | CRSCH | EMPLOYEE | 11/18/2023 10:22 AM | 11/13/2023 10:50 AM |  |  | SAH11225 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20490696SA | 440964 | GE Healthcare Technologies | 2016793-003 |  | ICU9 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1835599 | 11/14/2023 10:08 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-452-2477 / TEE Probe cleaner is not working - It is not doing what it was supposed to do to clean the probe. TM called the vendor, who advised to have CES call them at 919-255-9472. / Echo Dept. / | Crystal Schmitt | CRSCH | EMPLOYEE | 11/16/2023 10:08 AM | 11/27/2023 08:29 AM |  |  | SAH11284 | TEE Probe Disinfector | 26-654 | Disinfectors, Liquid Germicide, Transesophageal Echocardiography Probe | 20100164 | 456998 | CS Medical LLC | TCD100S8 |  | Echo Dept | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1835648 | 11/14/2023 12:53 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/19/2023 12:53 PM | 11/15/2023 09:36 AM |  |  | L41843 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84662 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1835650 | 11/14/2023 12:55 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to bottom housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/19/2023 12:55 PM | 11/15/2023 08:30 AM |  |  | L42150 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 068630 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1836373 | 11/16/2023 12:07 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Device not in range for set-point temperature | Crystal Schmitt | CRSCH | EMPLOYEE | 11/21/2023 12:07 PM | 11/17/2023 08:40 AM |  |  | L42276 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 076707 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1836391 | 11/16/2023 01:37 PM | CM | Corrective Maintenance | IP | In Progress | SHOP-STOCK orders November 2023 | Crystal Schmitt | CRSCH | EMPLOYEE | 11/21/2023 01:37 PM | 11/16/2023 01:37 PM |  |  | SHOP-STOCK | Shop Stock | CC-056 | Environment | Shop Stock | 000 | Misc-Numbers | Stock |  | SAH Hospital | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1836442 | 11/16/2023 06:55 PM | CM | Corrective Maintenance | WP | Awaiting Parts | CO2 Monitor is making a rattling noise. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/21/2023 06:55 PM | 11/22/2023 10:30 AM |  |  | SAH2694 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8530515 | 440964 | GE Healthcare Technologies | E-MiniC-00 | PACU | PACU | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH
1836574 | 11/17/2023 01:13 PM | CM | Corrective Maintenance | IP | In Progress | Device error indicating board failure. Rep on-site. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 01:13 PM | 11/17/2023 01:17 PM |  |  | SAH9152 | Electrosurgical Unit | 18-232 | Electrosurgical Units, Monopolar/Bipolar, Argon-Enhanced Coagulation | 11535108 | 108422 | ERBE USA Inc | 10160-100 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1836624 | 11/17/2023 03:20 PM | CM | Corrective Maintenance | IP | In Progress | Error message. Erbe tech onsite troubleshooting. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 03:20 PM | 11/17/2023 03:27 PM |  |  | SAH9153 | Argon Plasma Coag Module | 17-738 | Gas Delivery Units, Argon-Enhanced Coagulation | 11532372 | 108422 | ERBE USA Inc | APC3 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1836894 | 11/20/2023 09:37 AM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Not heating up to set point range. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/25/2023 09:37 AM | 11/22/2023 01:20 PM |  |  | L42271 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 106245 | 418683 | Arizant Healthcare Inc | 245 |  | Surgery | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1837670 | 11/27/2023 08:10 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Not sending gas flow | Crystal Schmitt | CRSCH | EMPLOYEE | 12/02/2023 08:10 AM | 11/27/2023 08:40 AM |  |  | SAH11120 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8805151 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | EC 24 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1837723 | 11/27/2023 10:38 AM | CM | Corrective Maintenance | OF | Out For Repair | PM Failed-Device not heating up to temp. Out of range for set point temperature. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/02/2023 10:38 AM | 11/27/2023 11:07 AM |  |  | L42239 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 91996 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1837835 | 11/27/2023 05:20 PM | CM | Corrective Maintenance | IP | In Progress | PM failed-Leak test failed | Crystal Schmitt | CRSCH | EMPLOYEE | 12/02/2023 05:20 PM | 11/27/2023 01:20 PM |  |  | SAH6271 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8529135 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | CVICU2 | 1105-13102 | SAH-Intensive Care Unit II 3NW | IU | Active | SAH | SAH