## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1835648 | 11/14/2023 12:53 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/19/2023 12:53 PM | 03/22/2024 01:20 PM |  |  | L41843 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84662 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1845993 | 12/08/2023 04:30 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Needs new level sensor3/4/24-Failed "inspect fluid delivery line" check for one valve | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 04:30 PM | 03/04/2024 01:34 PM |  |  | SAH6111 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYARY026 | 451492 | Medivance Inc A Bard Medical C | 5000 |  | ER- | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1857801 | 01/11/2024 09:12 AM | CM | Corrective Maintenance | IP | In Progress | RPT project OR | Crystal Schmitt | CRSCH | EMPLOYEE | 01/16/2024 09:12 AM | 03/19/2024 05:04 PM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1862213 | 01/29/2024 03:49 PM | CM | Corrective Maintenance | WP | Awaiting Parts | 727.502.4391 / PH Meter receives error during calibration process / Chemistry Dept, 1st Floor | Crystal Schmitt | CRSCH | EMPLOYEE | 02/03/2024 03:49 PM | 02/22/2024 03:22 PM |  |  | SAH11375 | Ph Meters | 15-164 | pH Meters | AE95014245 | 375242 | Fisher Healthcare Div Thermo S | AE 150 |  | SAH-Chemistry | 1105-31000 | SAH-Laboratory Services | 04 | Inventory | SAH | SAH
1870133 | 02/10/2024 05:48 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Two keypads are not workingLOC: Dialysis Water Room | Crystal Schmitt | CRSCH | EMPLOYEE | 02/15/2024 05:48 PM | 02/22/2024 09:15 AM |  |  | SAH1864 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 8T0S234256 | 312187 | Fresenius Medical Care North A | 2008 T |  | Dialysis Water Room | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1871138 | 02/15/2024 04:29 PM | CM | Corrective Maintenance | IP | In Progress | PB980 training | Crystal Schmitt | CRSCH | EMPLOYEE | 02/20/2024 04:29 PM | 02/15/2024 05:25 PM |  |  | TRAINING | Used for Training Work Orders | CC-056 | Environment | NSN | 000 | Misc-Numbers | Training |  |  | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1872905 | 02/20/2024 02:28 PM | CM | Corrective Maintenance | WP | Awaiting Parts | 727-825-1287 / Low pressure on the pump. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 02:28 PM | 03/13/2024 07:52 AM |  |  | SAH5288 | Endoscope Flushing Pump | 13-192 | Pumps | 21735979 | 276750 | Olympus America Inc | OFP-2 |  | Endo Procedure Rm 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1872943 | 02/20/2024 05:34 PM | CM | Corrective Maintenance | WP | Awaiting Parts | This device and other system components are plugged into a power transformer, which is plugged into a non-approved power strip. CM to correct power strip issue. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 05:34 PM | 02/20/2024 05:38 PM |  |  | SAH4002 | Stimulus Generator Unit | 17-898 | Monitoring Systems, Physiologic, Cardiac Electrophysiology | 4417-SGU/SC397 | 427969 | Bard Peripheral Vascular Div C | MP3008 |  | EP LAB | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1874209 | 02/26/2024 11:12 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118, Not passing negative pressure and tempLOC: Water Room, Dialysis | Crystal Schmitt | CRSCH | EMPLOYEE | 03/02/2024 11:12 AM | 02/27/2024 01:36 PM |  |  | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water Room, Dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1879712 | 03/05/2024 11:32 AM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Device failed leak check | Crystal Schmitt | CRSCH | EMPLOYEE | 03/10/2024 12:32 PM | 03/18/2024 03:38 PM |  |  | SAH5707 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8514910 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | EPACU04 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1880274 | 03/06/2024 07:52 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1008. Keeps reading low temperature. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/11/2024 08:52 AM | 03/06/2024 11:25 AM |  |  | SAH4037 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A867194 | 107875 | Exergen Corp | TAT-5000 | PACU | PACU | 1105-21500 | SAH-PACU | 04 | Inventory | SAH | SAH
1880788 | 03/08/2024 09:27 AM | CM | Corrective Maintenance | IP | In Progress | PDM not being recognized by B650 monitor, replacement OK | Crystal Schmitt | CRSCH | EMPLOYEE | 03/13/2024 10:27 AM | 03/11/2024 09:44 AM |  |  | SAH11215 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20430318SA | 440964 | GE Healthcare Technologies | 2016793-003 |  |  | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1882410 | 03/12/2024 10:55 AM | CM | Corrective Maintenance | IP | In Progress | Need key tether installed | Crystal Schmitt | CRSCH | EMPLOYEE | 03/17/2024 10:55 AM | 03/12/2024 10:59 AM |  |  | SAH5531 | Endo Motorized Travel Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 2000RFWDII082217A | 453029 | PHS West Inc | Ergo Express |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1883259 | 03/18/2024 03:10 AM | EN | Easy Net | IP | In Progress | T7RO with "insufficient water alarms" and "internal service alarms" Bay 3 HD suite portable hooked up to wall box at the time. No water issues when setting up the machine (rinsing, priming, testing etc.) but as soon as treatment started the RO wasn"t getting enough water. Machine tagged out in water room and Paul with biomed made aware. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/23/2024 03:10 AM | 03/21/2024 04:10 PM |  |  | SAH11396 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1397868 | 199732 | Mar Cor Purification A Cantel | WRO300H |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1883366 | 03/18/2024 11:52 AM | CM | Corrective Maintenance | WP | Awaiting Parts | petal isn''t working. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/23/2024 11:52 AM | 03/21/2024 08:20 AM |  |  | L49546 | Surgical Site Irrigators | 23-527 | Irrigators, Surgical Site | P02496 | 457057 | Medivators Corp A Cantel Medic | EGA-500 |  | Endo Storage Room 2 Traveling | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1883634 | 03/19/2024 08:03 PM | CM | Corrective Maintenance | IP | In Progress | Found power cord/charger damaged during department PMs | Crystal Schmitt | CRSCH | EMPLOYEE | 03/24/2024 08:03 PM | 03/19/2024 08:05 PM |  |  | SAH3370 | Printer | 22-203 | Printers, Computer/Network, Ink-Jet | U64418-G6W123430 | 473839 | Brother International | MW-260MFi |  | PIC TEAM | 1105-37200 | SAH-IV Therapy | 04 | Inventory | SAH | SAH
1885219 | 03/21/2024 12:44 PM | CM | Corrective Maintenance | IP | In Progress | PM failed-Failed pulse width, pulse frequency and stimulus duration test. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/26/2024 12:44 PM | 03/21/2024 04:00 PM |  |  | SAH5393 | Electroshock Unit | 11-484 | Stimulators, Electrical, Brain, Convulsive Therapy | 42685 | 104776 | Somatics LLC | System IV |  | SAH-PACU | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH
1885387 | 03/22/2024 11:01 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1125 or 727-692-9317/Not running blood pressure. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 11:01 AM | 03/22/2024 06:04 PM |  |  | SAH5881 | Front End Hemodynamic Monitor | 12-648 | Monitoring Systems, Physiologic, Cardiac Catheterization | US1550001478 | 452291 | Philips Interventional Patient | FC2020 |  | Cath Lab 2 | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1885408 | 03/22/2024 12:25 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Found during PM-Bottom corner LCD is dark and corner of front chassis damaged. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 12:25 PM | 03/22/2024 12:25 PM |  |  | SAH5687 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311035HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 09 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1885424 | 03/22/2024 01:17 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to top and bottom housing | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 01:17 PM | 03/22/2024 01:01 PM |  |  | L42214 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 82323 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH