## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1762073 | 05/10/2023 11:16 AM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Casing cracked. Needs to be sent out for repair. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/15/2023 11:16 AM | 05/25/2023 12:30 PM |  |  | L41840 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84659 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1762080 | 05/10/2023 11:31 AM | CM | Corrective Maintenance | OF | Out For Repair | PM Failed-Casing cracked | Crystal Schmitt | CRSCH | EMPLOYEE | 05/15/2023 11:31 AM | 06/09/2023 12:50 PM | PHY | Physical Damage | L41830 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 83861 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1764764 | 05/23/2023 04:18 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Bottom casing damaged. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/28/2023 04:18 PM | 05/25/2023 11:30 AM |  |  | L41841 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84658 | 418683 | Arizant Healthcare Inc | 245 |  | Loaned to OR room 4 (5/5/22) | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1776824 | 06/16/2023 11:15 AM | CM | Corrective Maintenance | OF | Out For Repair | Physical Damage-Found cracks in both the top and bottom housing. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/21/2023 11:15 AM | 06/19/2023 09:25 AM |  |  | L42123 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 56924 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1777608 | 06/21/2023 05:53 PM | CM | Corrective Maintenance | IP | In Progress | Vitals not crossing from bedside to cerner | Crystal Schmitt | CRSCH | EMPLOYEE | 06/26/2023 05:53 PM | 09/22/2023 07:24 AM |  |  | SAH8891 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | STF22232070SA | 440964 | GE Healthcare Technologies | B650 |  | SAH OR PUMP RM / AVANCE | 1105-20100 | SAH-Anesthesiology | IU | Active | SAH | SAH
1783204 | 07/07/2023 07:09 AM | CM | Corrective Maintenance | DU | Device Unavailable | PM Failed-Damage to housing | Crystal Schmitt | CRSCH | EMPLOYEE | 07/12/2023 07:09 AM | 09/11/2023 09:15 AM |  |  | SAH12480 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SQC19252128SA | 440964 | GE Healthcare Technologies | B650 |  | CVICU_03 | 1105-13102 | SAH-Intensive Care Unit II 3NW | IU | Active | SAH | SAH
1796507 | 08/05/2023 07:49 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118 - T5 - Machine has a leak-RO unit is leaking from the bottom. | Crystal Schmitt | CRSCH | EMPLOYEE | 08/10/2023 07:49 AM | 09/25/2023 09:38 AM |  |  | SAH2545 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1377811 | 199732 | Mar Cor Purification A Cantel | WRO300H |  | RO Machine #8 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1797533 | 08/10/2023 02:56 PM | CM | Corrective Maintenance | AD | Awaiting Delivery of Equipment | Capnography not functioning | Crystal Schmitt | CRSCH | EMPLOYEE | 08/15/2023 02:56 PM | 09/11/2023 10:25 AM |  |  | SAH8501 | Front End Hemodynamic Monitor | 12-648 | Monitoring Systems, Physiologic, Cardiac Catheterization | US1419000011 | 452291 | Philips Interventional Patient | FC2010 |  |  | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1798850 | 08/15/2023 02:16 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM Failed-Failed air delivery test at 35 SLPM out of MFG spec. | Crystal Schmitt | CRSCH | EMPLOYEE | 08/20/2023 02:16 PM | 08/21/2023 08:55 AM |  |  | SAH2916 | Non-Invasive Ventilator | 20-746 | Ventilators, Noninvasive Positive Pressure | 100365895 | 454161 | Philips Respironics | V60 |  | AH-V60-16 | 1105-35000 | SAH-Respiratory Care Services 2nd | LS | High Risk | SAH | SAH
1799545 | 08/21/2023 08:12 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1030 - filter was replaced over the weekend and wants to be sure it was done properly, there is something a little different - one item was changed but the other wasn''t (filters) | Crystal Schmitt | CRSCH | EMPLOYEE | 08/26/2023 08:12 AM | 09/27/2023 07:40 AM |  |  | SAH11283 | Disinfectant Neutralizer Stati | 26-862 | Neutralizing Stations, Disinfectant, Glutaraldehyde/Aldehyde-Based | CDS03-1017 | 456998 | CS Medical LLC | AC-DS-03 |  | B3530s | 1105-30620 | SAH-Non Invasive Cardiology | 04 | Inventory | SAH | SAH
1800353 | 08/24/2023 07:47 AM | CM | Corrective Maintenance | IP | In Progress | PM Failed-No ref gas flow | Crystal Schmitt | CRSCH | EMPLOYEE | 08/29/2023 07:47 AM | 09/13/2023 09:16 AM |  |  | SAH8982 | Airway Gas Module w/Agent Id | 20-778 | Physiologic Monitor Modules, Anesthetic Gas | SGV22100068HA | 440964 | GE Healthcare Technologies | E-sCAiO-00 |  |  | 1105-20100 | SAH-Anesthesiology | IU | Active | SAH | SAH
1800853 | 08/27/2023 02:04 AM | EN | Easy Net | OF | Out For Repair | Not reading when plugged into bedside monitor. | Crystal Schmitt | CRSCH | EMPLOYEE | 09/01/2023 02:04 AM | 09/25/2023 08:28 AM |  |  | SAH8854 | Bispectral Index BIS Module | 20-783 | Physiologic Monitor Modules, Electroencephalography | SGS22110069HA | 440964 | GE Healthcare Technologies | E-BIS-01 |  | SAH-ICU | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1801136 | 08/29/2023 08:52 AM | CM | Corrective Maintenance | OF | Out For Repair | Device left on broken shelf with note "The info on screen is not visible very well very faint" | Crystal Schmitt | CRSCH | EMPLOYEE | 09/03/2023 08:52 AM | 09/11/2023 10:24 AM |  |  | SAH4356 | Feeding Pump | 13-209 | Infusion Pumps, Enteral Feeding | C13051324 | 452294 | Covidien (US) | 382400 |  | MM | 1105-85160 | SAH-Sterile Processing | IU | Active | SAH | SAH
1806574 | 09/05/2023 02:27 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118 // SAH T3 // Dialysis - rotor plate came off. Still using machine | Crystal Schmitt | CRSCH | EMPLOYEE | 09/10/2023 02:27 PM | 09/20/2023 01:55 PM |  |  | SAH11906 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 2T0S264885 | 312187 | Fresenius Medical Care North A | 2008 T |  | T20 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1809239 | 09/18/2023 08:04 AM | CM | Corrective Maintenance | IP | In Progress | 727 502 4411 US machine, Sonisite - Error Code: P...... | Crystal Schmitt | CRSCH | EMPLOYEE | 09/20/2023 08:04 AM | 09/19/2023 02:01 PM |  |  | SAH8200 | Portable Ultrasound System | 18-143 | Scanning Systems, Ultrasonic, Portable | WK3D03 | 340440 | SonoSite div of FUJIFILM | P08778 |  | CATH LAB | 1105-30610 | SAH-Cardiac Cath Lab | 04 | Inventory | SAH | SAH
1809915 | 09/18/2023 12:46 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727.825.1118 / Problem: Not functioning at all - Low Flow Error - Low Temperature / LOC: Dialysis - T2 - Water Room | Crystal Schmitt | CRSCH | EMPLOYEE | 09/23/2023 12:46 PM | 09/18/2023 02:25 PM |  |  | SAH1864 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 8T0S234256 | 312187 | Fresenius Medical Care North A | 2008 T |  | T-6 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1810262 | 09/19/2023 02:49 PM | CM | Corrective Maintenance | IP | In Progress | 727-825-1118. Monitor will not turn on. | Crystal Schmitt | CRSCH | EMPLOYEE | 09/21/2023 02:49 PM | 09/25/2023 09:03 AM |  |  | SAH11441 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT22270048HA | 440964 | GE Healthcare Technologies | B450 |  | Dialysis Room 2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1810520 | 09/20/2023 04:50 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Spoke with Trinity Biomedical-Device needs a new board | Crystal Schmitt | CRSCH | EMPLOYEE | 09/25/2023 04:50 PM | 09/20/2023 05:04 PM |  |  | SAH4194 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248862 | 312187 | Fresenius Medical Care North A | 2008 T |  | Dialysis T15 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1810552 | 09/21/2023 07:14 AM | CM | Corrective Maintenance | IP | In Progress | 727825118- fails optical detector / b104-108 error after troubleshooting LOC: at2 | Crystal Schmitt | CRSCH | EMPLOYEE | 09/26/2023 07:14 AM | 09/26/2023 05:02 PM |  |  | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | at2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1811394 | 09/22/2023 04:38 PM | CM | Corrective Maintenance | IP | In Progress | Updating monitor alarm settings in ICU2 | Crystal Schmitt | CRSCH | EMPLOYEE | 09/27/2023 04:38 PM | 09/25/2023 08:54 AM |  |  | SPEC-PRO | Department Environment Rounds | CC-056 | Environment | Special Project | 000 | Misc-Numbers | ENVIRONMENT |  | SAH - Clinical Engineering | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1811481 | 09/24/2023 10:36 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1020, Broken latch on the pumpLOC: ICU 1 at the Nurse''s station | Crystal Schmitt | CRSCH | EMPLOYEE | 09/29/2023 10:36 AM | 09/25/2023 03:26 PM |  |  | SAH6623 | Feeding Pump | 13-209 | Infusion Pumps, Enteral Feeding | C17177985 | 452294 | Covidien (US) | 382400 |  | ICU 1 | 1105-85160 | SAH-Sterile Processing | IU | Active | SAH | SAH
1812100 | 09/25/2023 08:31 PM | CM | Corrective Maintenance | WP | Awaiting Parts | 727-825-1008Loc: PACU Main Desk Air Trap Cannister is not Registering and keeps prompting Errors | Crystal Schmitt | CRSCH | EMPLOYEE | 09/30/2023 08:31 PM | 09/26/2023 05:03 PM |  |  | SAH2706 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8529926 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | PACU Main Desk | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH
1812258 | 09/26/2023 02:59 PM | CM | Corrective Maintenance | IP | In Progress | 7278251287 - Error 302 - Endo room 2 | Crystal Schmitt | CRSCH | EMPLOYEE | 10/01/2023 02:59 PM | 09/26/2023 03:20 PM |  |  | SAH11427 | Video Processor, Endo | 18-034 | Image Processors, Video, Endoscopy | 7244402 | 276750 | Olympus America Inc | CV-190 |  | SAH-Endoscopy | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1812359 | 09/27/2023 07:44 AM | EN | Easy Net | IP | In Progress | not reading correctly. | Crystal Schmitt | CRSCH | EMPLOYEE | 10/02/2023 07:44 AM | 09/27/2023 02:02 PM |  |  | SAH9285 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A1383878 | 107875 | Exergen Corp | TAT-5000 |  | nursing station. | 1105-32100 | SAH-Hemodialysis | 04 | Inventory | SAH | SAH
1812376 | 09/27/2023 08:50 AM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | keep reading high and low error. Turn on and off but keep saying same thing. | Crystal Schmitt | CRSCH | EMPLOYEE | 10/02/2023 08:50 AM | 09/27/2023 02:00 PM |  |  | SAH6551 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 7T0S213143 | 312187 | Fresenius Medical Care North A | 2008 T |  | water room. | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH