## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1835648 | 11/14/2023 12:53 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/19/2023 12:53 PM | 02/26/2024 11:35 AM |  |  | L41843 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84662 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1845993 | 12/08/2023 04:30 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Needs new level sensor3/4/24-Failed "inspect fluid delivery line" check for one valve | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 04:30 PM | 03/04/2024 01:34 PM |  |  | SAH6111 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYARY026 | 451492 | Medivance Inc A Bard Medical C | 5000 |  | ER- | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1857801 | 01/11/2024 09:12 AM | CM | Corrective Maintenance | IP | In Progress | RPT project OR | Crystal Schmitt | CRSCH | EMPLOYEE | 01/16/2024 09:12 AM | 02/22/2024 11:30 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1862213 | 01/29/2024 03:49 PM | CM | Corrective Maintenance | WP | Awaiting Parts | 727.502.4391 / PH Meter receives error during calibration process / Chemistry Dept, 1st Floor | Crystal Schmitt | CRSCH | EMPLOYEE | 02/03/2024 03:49 PM | 02/22/2024 03:22 PM |  |  | SAH11375 | Ph Meters | 15-164 | pH Meters | AE95014245 | 375242 | Fisher Healthcare Div Thermo S | AE 150 |  | SAH-Chemistry | 1105-31000 | SAH-Laboratory Services | 04 | Inventory | SAH | SAH
1869476 | 02/07/2024 11:17 AM | CM | Corrective Maintenance | AF | Awaiting FSR | Software version does not match other telemetry servers. Issue with power distribution to servers and receivers in MDF closet. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/12/2024 11:17 AM | 02/20/2024 11:05 AM |  |  | SAH3336 | Telemetry Server | 21-973 | Network Servers | SEE09100133GA | 440964 | GE Healthcare Technologies | MP100R |  | MDF-AT2 | 1105-18415 | SAH-Neuroscience Unit 5W | IU | Active | SAH | SAH
1870133 | 02/10/2024 05:48 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Two keypads are not workingLOC: Dialysis Water Room | Crystal Schmitt | CRSCH | EMPLOYEE | 02/15/2024 05:48 PM | 02/22/2024 09:15 AM |  |  | SAH1864 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 8T0S234256 | 312187 | Fresenius Medical Care North A | 2008 T |  | Dialysis Water Room | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1871138 | 02/15/2024 04:29 PM | CM | Corrective Maintenance | IP | In Progress | PB980 training | Crystal Schmitt | CRSCH | EMPLOYEE | 02/20/2024 04:29 PM | 02/15/2024 05:25 PM |  |  | TRAINING | Used for Training Work Orders | CC-056 | Environment | NSN | 000 | Misc-Numbers | Training |  |  | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1872905 | 02/20/2024 02:28 PM | CM | Corrective Maintenance | WP | Awaiting Parts | 727-825-1287 / Low pressure on the pump. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 02:28 PM | 03/13/2024 07:52 AM |  |  | SAH5288 | Endoscope Flushing Pump | 13-192 | Pumps | 21735979 | 276750 | Olympus America Inc | OFP-2 |  | Endo Procedure Rm 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1872943 | 02/20/2024 05:34 PM | CM | Corrective Maintenance | WP | Awaiting Parts | This device and other system components are plugged into a power transformer, which is plugged into a non-approved power strip. CM to correct power strip issue. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 05:34 PM | 02/20/2024 05:38 PM |  |  | SAH4002 | Stimulus Generator Unit | 17-898 | Monitoring Systems, Physiologic, Cardiac Electrophysiology | 4417-SGU/SC397 | 427969 | Bard Peripheral Vascular Div C | MP3008 |  | EP LAB | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1874209 | 02/26/2024 11:12 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118, Not passing negative pressure and tempLOC: Water Room, Dialysis | Crystal Schmitt | CRSCH | EMPLOYEE | 03/02/2024 11:12 AM | 02/27/2024 01:36 PM |  |  | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water Room, Dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1874232 | 02/26/2024 12:10 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damaged housing; both top and bottom plastic cracked. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/02/2024 12:10 PM | 03/04/2024 12:04 PM |  |  | L41902 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 87724 | 418683 | Arizant Healthcare Inc | 245 |  | SAH-Surgery | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1875526 | 02/29/2024 03:59 PM | CM | Corrective Maintenance | IP | In Progress | Won''t stop running - B3530s - 727-825-1030 x171471 | Crystal Schmitt | CRSCH | EMPLOYEE | 03/05/2024 03:59 PM | 02/29/2024 04:27 PM |  |  | SAH11283 | Disinfectant Neutralizer Stati | 26-862 | Neutralizing Stations, Disinfectant, Glutaraldehyde/Aldehyde-Based | CDS03-1017 | 456998 | CS Medical LLC | AC-DS-03 |  | B3530s | 1105-30620 | SAH-Non Invasive Cardiology | 04 | Inventory | SAH | SAH
1879280 | 03/03/2024 08:28 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1055 / EKG machine not turning on / 5W Nursing Station Electrical Room | Crystal Schmitt | CRSCH | EMPLOYEE | 03/08/2024 08:28 AM | 03/07/2024 08:10 AM |  |  | SAH3176 | EKG Mulitchannel Interp | 16-231 | Electrocardiographs, Multichannel, Interpretive | SGM11450651PA | 440964 | GE Healthcare Technologies | MAC 5500HD |  | 5W Nursing Station | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1879712 | 03/05/2024 11:32 AM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Device failed leak check | Crystal Schmitt | CRSCH | EMPLOYEE | 03/10/2024 12:32 PM | 03/05/2024 11:34 AM |  |  | SAH5707 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8514910 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | EPACU04 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1880182 | 03/05/2024 01:02 PM | CM | Corrective Maintenance | IP | In Progress | Documenting return of loaner device | Crystal Schmitt | CRSCH | EMPLOYEE | 03/10/2024 02:02 PM | 03/13/2024 04:20 PM |  |  | R31325 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYERY004 | 451492 | Medivance Inc A Bard Medical C | 5000 |  |  | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1880249 | 03/05/2024 06:38 PM | CM | Corrective Maintenance | IP | In Progress | All ACU monitors printing to CVICU printer | Crystal Schmitt | CRSCH | EMPLOYEE | 03/10/2024 07:38 PM | 03/05/2024 06:41 PM |  |  | SAH3088 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNE18502036SA | 440964 | GE Healthcare Technologies | B450 |  | SAH-Amb. Care-Same Day Rm 9 | 1105-21300 | SAH-Ambulatory Care Unit | IU | Active | SAH | SAH
1880274 | 03/06/2024 07:52 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1008. Keeps reading low temperature. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/11/2024 08:52 AM | 03/06/2024 11:25 AM |  |  | SAH4037 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A867194 | 107875 | Exergen Corp | TAT-5000 | PACU | PACU | 1105-21500 | SAH-PACU | 04 | Inventory | SAH | SAH
1880788 | 03/08/2024 09:27 AM | CM | Corrective Maintenance | IP | In Progress | PDM not being recognized by B650 monitor, replacement OK | Crystal Schmitt | CRSCH | EMPLOYEE | 03/13/2024 10:27 AM | 03/11/2024 09:44 AM |  |  | SAH11215 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20430318SA | 440964 | GE Healthcare Technologies | 2016793-003 |  |  | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1882410 | 03/12/2024 10:55 AM | CM | Corrective Maintenance | IP | In Progress | Need key tether installed | Crystal Schmitt | CRSCH | EMPLOYEE | 03/17/2024 10:55 AM | 03/12/2024 10:59 AM |  |  | SAH5531 | Endo Motorized Travel Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 2000RFWDII082217A | 453029 | PHS West Inc | Ergo Express |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1882806 | 03/13/2024 07:27 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Leak test failed | Crystal Schmitt | CRSCH | EMPLOYEE | 03/18/2024 07:27 PM | 03/13/2024 07:28 PM |  |  | SAH7845 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8519140 | 440964 | GE Healthcare Technologies | E-MiniC-00 | PACU | PACU | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH
1882807 | 03/13/2024 07:30 PM | CM | Corrective Maintenance | IP | In Progress | PM failed-Leak check failed | Crystal Schmitt | CRSCH | EMPLOYEE | 03/18/2024 07:30 PM | 03/13/2024 08:10 PM |  |  | SAH11137 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8805135 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | EC 13 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH