## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1823628 | 10/24/2023 01:50 PM | CM | Corrective Maintenance | OF | Out For Repair | Reading incorrect ground resistance | Crystal Schmitt | CRSCH | EMPLOYEE | 10/29/2023 01:50 PM | 10/26/2023 09:20 AM |  |  | SAH8877 | Electrical Safety Analyzer | CC-040 | Test Equipment | 5942040 | 178848 | Fluke Biomedical | ESA609 |  |  | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1829136 | 11/02/2023 12:10 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed due to physical damage, needs new front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/07/2023 11:10 AM | 11/15/2023 12:20 PM |  |  | SAH3229 | LP20e Defibrillator/Pacer | 17-882B | Defibrillator/Pacemakers, External, Biphasic | 48234721 | 452575 | Physio-Control Inc | LP20E |  | SAH-3rd EP lab | 1105-13100 | SAH-Intensive Care Unit | LS | High Risk | SAH | SAH
1835395 | 11/13/2023 10:22 AM | CM | Corrective Maintenance | OF | Out For Repair | Device says "Reconnect PDM" error message | Crystal Schmitt | CRSCH | EMPLOYEE | 11/18/2023 10:22 AM | 12/01/2023 10:00 AM |  |  | SAH11225 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20490696SA | 440964 | GE Healthcare Technologies | 2016793-003 |  | ICU9 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1835648 | 11/14/2023 12:53 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/19/2023 12:53 PM | 11/15/2023 09:36 AM |  |  | L41843 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84662 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1836456 | 11/17/2023 06:33 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 7278251118 - error high/low flow, temp & negative pressure | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 06:33 AM | 12/19/2023 02:45 PM | COM | Component Failure-NOT Battery | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water RM AT 2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1836574 | 11/17/2023 01:13 PM | CM | Corrective Maintenance | OF | Out For Repair | Device error indicating board failure. Rep on-site. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 01:13 PM | 11/29/2023 09:01 AM |  |  | SAH9152 | Electrosurgical Unit | 18-232 | Electrosurgical Units, Monopolar/Bipolar, Argon-Enhanced Coagulation | 11535108 | 108422 | ERBE USA Inc | 10160-100 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1836624 | 11/17/2023 03:20 PM | CM | Corrective Maintenance | OF | Out For Repair | Error message. Erbe tech onsite troubleshooting. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 03:20 PM | 11/29/2023 09:06 AM |  |  | SAH9153 | Argon Plasma Coag Module | 17-738 | Gas Delivery Units, Argon-Enhanced Coagulation | 11532372 | 108422 | ERBE USA Inc | APC3 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1837025 | 11/20/2023 06:12 PM | CM | Corrective Maintenance | OF | Out For Repair | Issue: Device is giving low flow error.LOC: SAH CVICU.Contact: 7278251039 | Crystal Schmitt | CRSCH | EMPLOYEE | 11/25/2023 06:12 PM | 12/13/2023 01:25 PM |  |  | SAH9217 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYGZAL029 | 451492 | Medivance Inc A Bard Medical C | 6000 | CVICU | cvICU | 1105-12530 | SAH-Surgical Unit I AT4 | IU | Active | SAH | SAH
1838000 | 11/28/2023 02:05 PM | EN | Easy Net | WP | Awaiting Parts | Manoscan machine in ENDO preop room 5, has been failing to pressurize during calibration of probe. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/03/2023 02:05 PM | 11/29/2023 05:13 PM |  |  | SAH5651 | ManoScan Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 6071 | 391192 | Medtronic Given Imaging Inc | 8039 |  | Endoscopy, preop room 5 | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1840978 | 12/01/2023 07:52 PM | EN | Easy Net | OF | Out For Repair | only reads 85.5 for every patient...needs repair. Thanks! | Crystal Schmitt | CRSCH | EMPLOYEE | 12/06/2023 07:52 PM | 12/18/2023 03:08 PM |  |  | SAH4009 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A861278 | 107875 | Exergen Corp | TAT-5000 |  | SAH ICU-1 | 1105-13100 | SAH-Intensive Care Unit | 04 | Inventory | SAH | SAH
1841411 | 12/04/2023 09:35 AM | EN | Easy Net | OF | Out For Repair | Bladder scanner not sensing any urine in the bladder. Tried to scan with a different scanner that worked just fine. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/09/2023 09:35 AM | 12/07/2023 12:00 PM |  |  | SAH9191 | Bladder Vol and Mass Ultrasnd | 34-698 | Scanning Systems, Ultrasonic, Small-Parts, Urinary Bladder | C1520889 | 451928 | Verathon Inc | Bladderscan Prime PL |  | Electrical Closet | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1845848 | 12/08/2023 10:20 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Circuit device is plugged into is getting overpowered. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 10:20 AM | 12/11/2023 08:40 AM |  |  | SAH8387 | Mobile C-Arm | 11-758 | Radiographic/Fluoroscopic Units, Mobile (C-ARM) | 52011 | 399199 | Siemens Medical Solutions USA | 10308194 |  | SAH2 Endo | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1845993 | 12/08/2023 04:30 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Needs new level sensor | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 04:30 PM | 12/11/2023 08:07 AM |  |  | SAH6111 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYARY026 | 451492 | Medivance Inc A Bard Medical C | 5000 |  | ER- | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1846457 | 12/12/2023 10:42 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1042 - SAH1878 - Does not turn on - Have tried powering on, and still nothing. Location: SAH - Stress Department - 3rd Floor | Crystal Schmitt | CRSCH | EMPLOYEE | 12/17/2023 10:42 AM | 12/13/2023 02:15 PM |  |  | SAH1878 | Stress Excercise System | 17-723 | Monitoring Systems, Physiologic, Stress Exercise, Cardiac | SHS12331136SA | 440964 | GE Healthcare Technologies | 2012492-011 |  | Treadmill Rm 3rd Fl | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1846703 | 12/13/2023 08:37 AM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | Signature suite; cannot select different scenery. Reset the signature suite generator in the Steris Rack without resolution. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/18/2023 08:37 AM | 12/18/2023 07:42 AM |  |  | SAH11475 | OR Video Integration Systems | 20-388 | Automation Systems, Operating Room | R221790458 | 107114 | STERIS Corp | HexaVue System |  | OR 17 | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1846871 | 12/13/2023 03:21 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Device turned off and will not power back on | Crystal Schmitt | CRSCH | EMPLOYEE | 12/18/2023 03:21 PM | 12/14/2023 10:50 AM |  |  | SAH8863 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1407677 | 199732 | Mar Cor Purification A Cantel | MILLENIUM HX |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1849656 | 12/17/2023 10:35 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1020, Temperature probe is not reading correctly, message says low.LOC:ICU 2nd Floor | Crystal Schmitt | CRSCH | EMPLOYEE | 12/22/2023 10:35 AM | 12/20/2023 10:45 AM | 210 | Could not duplicate | SAH5322 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | 2928 | 451492 | Medivance Inc A Bard Medical C | 5000 |  | ICU 2nd Floor | 1105-13102 | SAH-Intensive Care Unit II 3NW | IU | Active | SAH | SAH
1849963 | 12/18/2023 12:34 PM | CM | Corrective Maintenance | IP | In Progress | Damaged housing | Crystal Schmitt | CRSCH | EMPLOYEE | 12/23/2023 12:34 PM | 12/20/2023 08:30 AM |  |  | L42212 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 103969 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1852282 | 12/28/2023 05:59 AM | CM | Corrective Maintenance | OP | Open | 727-825-1118 Loc: Equipment Repair Room - Dialysis Unable to Bleach - Error no Chemical Uptake | Crystal Schmitt | CRSCH | EMPLOYEE | 01/02/2024 05:59 AM | 12/28/2023 06:00 AM |  |  | SAH6551 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 7T0S213143 | 312187 | Fresenius Medical Care North A | 2008 T |  | TRAVEL #2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH