## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1823628 | 10/24/2023 01:50 PM | CM | Corrective Maintenance | OF | Out For Repair | Reading incorrect ground resistance | Crystal Schmitt | CRSCH | EMPLOYEE | 10/29/2023 01:50 PM | 10/26/2023 09:20 AM |  |  | SAH8877 | Electrical Safety Analyzer | CC-040 | Test Equipment | 5942040 | 178848 | Fluke Biomedical | ESA609 |  |  | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1829136 | 11/02/2023 12:10 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed due to physical damage, needs new front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/07/2023 11:10 AM | 11/15/2023 12:20 PM |  |  | SAH3229 | LP20e Defibrillator/Pacer | 17-882B | Defibrillator/Pacemakers, External, Biphasic | 48234721 | 452575 | Physio-Control Inc | LP20E |  | SAH-3rd EP lab | 1105-13100 | SAH-Intensive Care Unit | LS | High Risk | SAH | SAH
1835648 | 11/14/2023 12:53 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/19/2023 12:53 PM | 11/15/2023 09:36 AM |  |  | L41843 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84662 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1836456 | 11/17/2023 06:33 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 7278251118 - error high/low flow, temp & negative pressure | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 06:33 AM | 12/19/2023 02:45 PM | COM | Component Failure-NOT Battery | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water RM AT 2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1836574 | 11/17/2023 01:13 PM | CM | Corrective Maintenance | OF | Out For Repair | Device error indicating board failure. Rep on-site. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 01:13 PM | 11/29/2023 09:01 AM |  |  | SAH9152 | Electrosurgical Unit | 18-232 | Electrosurgical Units, Monopolar/Bipolar, Argon-Enhanced Coagulation | 11535108 | 108422 | ERBE USA Inc | 10160-100 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1836624 | 11/17/2023 03:20 PM | CM | Corrective Maintenance | OF | Out For Repair | Error message. Erbe tech onsite troubleshooting. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 03:20 PM | 11/29/2023 09:06 AM |  |  | SAH9153 | Argon Plasma Coag Module | 17-738 | Gas Delivery Units, Argon-Enhanced Coagulation | 11532372 | 108422 | ERBE USA Inc | APC3 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1837025 | 11/20/2023 06:12 PM | CM | Corrective Maintenance | OF | Out For Repair | Issue: Device is giving low flow error.LOC: SAH CVICU.Contact: 7278251039 | Crystal Schmitt | CRSCH | EMPLOYEE | 11/25/2023 06:12 PM | 01/08/2024 02:30 PM |  |  | SAH9217 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYGZAL029 | 451492 | Medivance Inc A Bard Medical C | 6000 | CVICU | cvICU | 1105-12530 | SAH-Surgical Unit I AT4 | IU | Active | SAH | SAH
1838000 | 11/28/2023 02:05 PM | EN | Easy Net | WP | Awaiting Parts | Manoscan machine in ENDO preop room 5, has been failing to pressurize during calibration of probe. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/03/2023 02:05 PM | 12/29/2023 09:57 AM |  |  | SAH5651 | ManoScan Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 6071 | 391192 | Medtronic Given Imaging Inc | 8039 |  | Endoscopy, preop room 5 | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1841411 | 12/04/2023 09:35 AM | EN | Easy Net | OF | Out For Repair | Bladder scanner not sensing any urine in the bladder. Tried to scan with a different scanner that worked just fine. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/09/2023 09:35 AM | 12/07/2023 12:00 PM |  |  | SAH9191 | Bladder Vol and Mass Ultrasnd | 34-698 | Scanning Systems, Ultrasonic, Small-Parts, Urinary Bladder | C1520889 | 451928 | Verathon Inc | Bladderscan Prime PL |  | Electrical Closet | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1845848 | 12/08/2023 10:20 AM | CM | Corrective Maintenance | AF | Awaiting FSR | Circuit device is plugged into is getting overpowered. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 10:20 AM | 01/04/2024 11:50 AM |  |  | SAH8387 | Mobile C-Arm | 11-758 | Radiographic/Fluoroscopic Units, Mobile (C-ARM) | 52011 | 399199 | Siemens Medical Solutions USA | 10308194 |  | SAH2 Endo | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1845993 | 12/08/2023 04:30 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Needs new level sensor | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 04:30 PM | 01/02/2024 09:24 AM |  |  | SAH6111 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYARY026 | 451492 | Medivance Inc A Bard Medical C | 5000 |  | ER- | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1846457 | 12/12/2023 10:42 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1042 - SAH1878 - Does not turn on - Have tried powering on, and still nothing. Location: SAH - Stress Department - 3rd Floor | Crystal Schmitt | CRSCH | EMPLOYEE | 12/17/2023 10:42 AM | 12/13/2023 02:15 PM |  |  | SAH1878 | Stress Excercise System | 17-723 | Monitoring Systems, Physiologic, Stress Exercise, Cardiac | SHS12331136SA | 440964 | GE Healthcare Technologies | 2012492-011 |  | Treadmill Rm 3rd Fl | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1846703 | 12/13/2023 08:37 AM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | Signature suite; cannot select different scenery. Reset the signature suite generator in the Steris Rack without resolution. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/18/2023 08:37 AM | 01/09/2024 09:01 AM |  |  | SAH11475 | OR Video Integration Systems | 20-388 | Automation Systems, Operating Room | R221790458 | 107114 | STERIS Corp | HexaVue System |  | OR 17 | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1846871 | 12/13/2023 03:21 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Device turned off and will not power back on | Crystal Schmitt | CRSCH | EMPLOYEE | 12/18/2023 03:21 PM | 01/03/2024 10:12 AM |  |  | SAH8863 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1407677 | 199732 | Mar Cor Purification A Cantel | MILLENIUM HX |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1852282 | 12/28/2023 05:59 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118Loc: Equipment Repair Room - Dialysis Unable to Bleach - Error no Chemical Uptake | Crystal Schmitt | CRSCH | EMPLOYEE | 01/02/2024 05:59 AM | 12/28/2023 03:16 PM |  |  | SAH6551 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 7T0S213143 | 312187 | Fresenius Medical Care North A | 2008 T |  | TRAVEL #2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1852470 | 12/28/2023 04:51 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PX 727-825-1232, Needs to be looked at, unit is not working. TM said it can be down first then tomorrow morning, but requests to have it STAT, due to only having on working unit.LOC: SAH1 Rm10 | Crystal Schmitt | CRSCH | EMPLOYEE | 12/30/2023 04:51 PM | 01/09/2024 02:34 PM |  |  | SAH8326 | Portable Ultrasound System | 18-143 | Scanning Systems, Ultrasonic, Portable | 052QNG | 340440 | SonoSite div of FUJIFILM | Sonosite PX |  | SAH1 Rm10 | 1105-34110 | SAH-Angiocardiography | 04 | Inventory | SAH | SAH
1853577 | 01/02/2024 08:26 AM | CM | Corrective Maintenance | IP | In Progress | Special Projects and Assignments January 2024 | Crystal Schmitt | CRSCH | EMPLOYEE | 01/07/2024 08:26 AM | 01/09/2024 08:25 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1853584 | 01/02/2024 08:52 AM | EN | Easy Net | OF | Out For Repair | thermometer case cracked, in Kelly"s office | Crystal Schmitt | CRSCH | EMPLOYEE | 01/07/2024 08:52 AM | 01/04/2024 10:10 AM |  |  | SAH1697 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A806867 | 107875 | Exergen Corp | TAT-5000 | PACU | PACU | 1105-21500 | SAH-PACU | 04 | Inventory | SAH | SAH
1856483 | 01/04/2024 04:08 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118 / Error message Diasafe fail is occurring on device | Crystal Schmitt | CRSCH | EMPLOYEE | 01/09/2024 04:08 AM | 01/05/2024 08:30 AM |  |  | SAH6551 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 7T0S213143 | 312187 | Fresenius Medical Care North A | 2008 T |  | repair room - New tower | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1856674 | 01/04/2024 04:28 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM Failed-Component loose inside of device | Crystal Schmitt | CRSCH | EMPLOYEE | 01/09/2024 04:28 PM | 01/05/2024 09:10 AM |  |  | SAH8533 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20190431SA | 440964 | GE Healthcare Technologies | 2016793-003 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1856930 | 01/06/2024 09:08 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-551-6849Location: HemodialysisKeeps saying: 104108 stuck closed | Crystal Schmitt | CRSCH | EMPLOYEE | 01/11/2024 09:08 AM | 01/08/2024 02:47 PM |  |  | SAH4188 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248874 | 312187 | Fresenius Medical Care North A | 2008 T |  | T10 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1857092 | 01/06/2024 10:05 PM | EN | Easy Net | IP | In Progress | Doppler is not working properly. | Crystal Schmitt | CRSCH | EMPLOYEE | 01/11/2024 10:05 PM | 01/09/2024 03:48 PM |  |  | SAH5352 | US Doppler Flow Detector | 10-429 | Detectors, Blood Flow, Ultrasonic | 811B-800002457 | 102998 | Parks Medical Electronics Inc | 811BL | ICU | ICU | 1105-13100 | SAH-Intensive Care Unit | 04 | Inventory | SAH | SAH
1857164 | 01/08/2024 07:04 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118 / No diasafe filter, need it installed. | Crystal Schmitt | CRSCH | EMPLOYEE | 01/13/2024 07:04 AM | 01/09/2024 08:52 AM |  |  | SAH6551 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 7T0S213143 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water Room | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1857408 | 01/09/2024 10:13 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Problem: V105 stock conductivity test failed.LOC: 2nd Fl - DialysisPhone: 727.825.1118 | Crystal Schmitt | CRSCH | EMPLOYEE | 01/14/2024 10:13 AM | 01/09/2024 02:21 PM |  |  | SAH4194 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248862 | 312187 | Fresenius Medical Care North A | 2008 T |  | Dialysis T15 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1857424 | 01/09/2024 11:06 AM | CM | Corrective Maintenance | IP | In Progress | Broken screen | Crystal Schmitt | CRSCH | EMPLOYEE | 01/14/2024 11:06 AM | 01/09/2024 03:41 PM |  |  | SAH9318 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A1386483 | 107875 | Exergen Corp | TAT-5000 |  | ENDOSCOPY | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1857481 | 01/09/2024 02:34 PM | CM | Corrective Maintenance | WP | Awaiting Parts | 727-452-2477 ¦ Treadmill (may be tagged wrong) --> Updated Tag Number SAH11282 need wire to connect treadmill to echo machine (needs replacing before the 16th) | Crystal Schmitt | CRSCH | EMPLOYEE | 01/14/2024 02:34 PM | 01/09/2024 03:27 PM |  |  | SAH11282 | Stress Exercise System | 17-723 | Monitoring Systems, Physiologic, Stress Exercise, Cardiac | STE21220011SA | 440964 | GE Healthcare Technologies | CASE V6.7 |  |  | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1857522 | 01/09/2024 06:17 PM | CM | Corrective Maintenance | OP | Open | 727-825-1020 / Making strange, loud noises. | Crystal Schmitt | CRSCH | EMPLOYEE | 01/14/2024 06:17 PM | 01/09/2024 06:17 PM |  |  | SAH9225 | US Doppler Flow Detector | 10-429 | Detectors, Blood Flow, Ultrasonic | 811B-800004365 | 102998 | Parks Medical Electronics Inc | 811BL |  | ICU 1 charge desk | 1105-13100 | SAH-Intensive Care Unit | 04 | Inventory | SAH | SAH