## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1872943 | 02/20/2024 05:34 PM | CM | Corrective Maintenance | WP | Awaiting Parts | This device and other system components are plugged into a power transformer, which is plugged into a non-approved power strip. CM to correct power strip issue. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 05:34 PM | 06/24/2024 12:00 PM |  |  | SAH4002 | Stimulus Generator Unit | 17-898 | Monitoring Systems, Physiologic, Cardiac Electrophysiology | 4417-SGU/SC397 | 427969 | Bard Peripheral Vascular Div C | MP3008 |  | EP LAB | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1892294 | 04/04/2024 06:26 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM Failed-USB port damaged-unable to block out port or connect any USB. Bottom corner of display is dark-display needs to be replaced. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/09/2024 06:26 PM | 07/15/2024 11:59 AM |  |  | SAH8034 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417321139HA | 440964 | GE Healthcare Technologies | B650 |  | EPACU01 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1892711 | 04/08/2024 11:42 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Found during PM-Damaged front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 04/13/2024 11:42 AM | 06/24/2024 01:58 PM |  |  | SAH5608 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311028HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 16 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1906860 | 05/16/2024 08:56 AM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Set point temp out of range. Only reaches 39.3C. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/21/2024 08:56 AM | 05/16/2024 11:01 AM |  |  | L42239 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 91996 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1911478 | 05/17/2024 11:43 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Found during PM-Caster breaks not functional. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/22/2024 11:43 AM | 07/12/2024 05:12 PM |  |  | SAH6364 | Mobile Exam Light | 33-257 | Lights, Examination, Stand-Alone, High End | 0819637 | 107071 | Medical Illumination Internat | 061523 |  |  | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1912625 | 05/25/2024 10:24 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Device reading "low voltage". Power management board issue | Crystal Schmitt | CRSCH | EMPLOYEE | 05/30/2024 10:24 AM | 05/28/2024 10:30 AM |  |  | SAH11072 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT20492117SA | 440964 | GE Healthcare Technologies | b450 V3 |  | EX06 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1913056 | 05/28/2024 11:25 AM | EN | Easy Net | IP | In Progress | when scanning patient it states there is no order for patient Note: This device is being relocated from 6N to Cardiac Medical II-3E once the repairs are completed. JM | Crystal Schmitt | CRSCH | EMPLOYEE | 06/02/2024 11:25 AM | 05/28/2024 06:59 PM |  |  | SAH1885 | EKG Mulitchannel Interp | 16-231 | Electrocardiographs, Multichannel, Interpretive | SKJ17210068PA | 440964 | GE Healthcare Technologies | MAC 5500HD |  | 6 N / LOC 901 | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1913215 | 05/28/2024 03:59 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damaged housing | Crystal Schmitt | CRSCH | EMPLOYEE | 06/02/2024 03:59 PM | 07/16/2024 08:30 AM |  |  | SAH9805 | Laparoscopic Insufflator | 16-849 | Insufflators, Laparoscopic | C03601 | 457057 | Medivators Corp A Cantel Medic | EGA-501 | ENDO | Endo 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1920452 | 06/06/2024 03:02 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Found during Steris PM-Plastic cover needs replaced | Crystal Schmitt | CRSCH | EMPLOYEE | 06/11/2024 03:02 PM | 07/16/2024 02:39 PM |  |  | SAH5342 | Surgical Boom | 22-613 | Facility Booms, Ceiling-Mounted | 0420017008 | 107114 | STERIS Corp | HarmonyAIR |  | Endo Rm 3 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1923763 | 06/20/2024 08:31 AM | EN | Easy Net | OF | Out For Repair | Myron not calibrating. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/25/2024 08:31 AM | 07/10/2024 12:30 PM |  |  | SAH7017 | Dialysate Conductivity Meters | 11-206 | Electrical Conductivity Meters, Dialysate | 497166 | 101527 | Myron L Co | D-1 |  | HEMODIALYSIS LAB | 1105-32100 | SAH-Hemodialysis | 04 | Inventory | SAH | SAH
1923794 | 06/20/2024 10:05 AM | CM | Corrective Maintenance | IP | In Progress | Error USB device not recognized. Vendor ticket number is mst0632832. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/25/2024 10:05 AM | 07/10/2024 07:55 AM |  |  | SAH9608 | Data Recorder Kit | 34-626 | Recorders, Physiologic, pH, Upper Gastrointestinal Tract, Ingestible Capsule | DR30720763 | 391192 | Medtronic Given Imaging Inc | FGS-0347 |  | Endoscopy Phys Dictation | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1924425 | 06/24/2024 01:47 PM | CM | Corrective Maintenance | IP | In Progress | Leakage tester used with Tee clean needs to be sent for calibration and added to AIMs. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/29/2024 01:47 PM | 06/24/2024 09:30 AM |  |  | ENV-EKG | Department Environment Rounds | CC-056 | Environment | Environment | 000 | Misc-Numbers | ENVIRONMENT |  | SAH-EKG-Echo Dept | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1925277 | 06/28/2024 02:42 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damaged rear housing and front housing-needs to go in for depot repair. | Crystal Schmitt | CRSCH | EMPLOYEE | 07/03/2024 02:42 PM | 07/02/2024 04:15 PM |  |  | SAH11749 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNE19262042SA | 440964 | GE Healthcare Technologies | B450 |  | TRANSPORT | 1105-13102 | SAH-Intensive Care Unit II 3NW | IU | Active | SAH | SAH
1925300 | 06/28/2024 05:51 PM | CM | Corrective Maintenance | IP | In Progress | Documenting adding Aeroscout tags to portable monitors in Endo prep rooms. | Crystal Schmitt | CRSCH | EMPLOYEE | 07/03/2024 05:51 PM | 06/28/2024 06:06 PM |  |  | ENV-ENDO | Department Environment Rounds | CC-056 | Environment | Environment | 000 | Misc-Numbers | ENVIRONMENT |  | SAH-Endoscopy | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1930708 | 07/04/2024 08:04 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118 ¦ Conductivity stuck at 35. ¦ SAH - Water Room 2nd Floor | Crystal Schmitt | CRSCH | EMPLOYEE | 07/09/2024 08:04 AM | 07/10/2024 10:40 AM |  |  | SAH4489 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1397865 | 199732 | Mar Cor Purification A Cantel | WRO300H |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1931600 | 07/08/2024 05:13 PM | EN | Easy Net | IP | In Progress | does not read the CO2, says there is an error and makes a whinning noise | Crystal Schmitt | CRSCH | EMPLOYEE | 07/13/2024 05:13 PM | 07/10/2024 05:16 PM |  |  | SAH7844 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8519121 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | PACU 22 | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH
1932626 | 07/15/2024 08:05 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Top case needs to be replaced. Needs battery and battery door. | Crystal Schmitt | CRSCH | EMPLOYEE | 07/20/2024 08:05 AM | 07/15/2024 08:37 AM |  |  | SAH11210 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20490750SA | 440964 | GE Healthcare Technologies | 2016793-003 |  |  | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1932830 | 07/16/2024 09:17 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Ordering spare B650 power cords | Crystal Schmitt | CRSCH | EMPLOYEE | 07/21/2024 09:17 AM | 07/16/2024 10:36 AM |  |  | SHOP-STOCK | Shop Stock | CC-056 | Environment | Shop Stock | 000 | Misc-Numbers | Stock |  | SAH Hospital | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1932997 | 07/16/2024 12:51 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118 ¦ No Water Alarm is Going Off but Water is in room ¦ AT2 Dialysis Arrow Room | Crystal Schmitt | CRSCH | EMPLOYEE | 07/21/2024 12:51 PM | 07/16/2024 04:05 PM |  |  | SAH1864 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 8T0S234256 | 312187 | Fresenius Medical Care North A | 2008 T |  | T-6 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH