## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1809239 | 09/18/2023 08:04 AM | CM | Corrective Maintenance | IP | In Progress | 727 502 4411 US machine, Sonisite - Error Code: P...... | Crystal Schmitt | CRSCH | EMPLOYEE | 09/23/2023 08:04 AM | 10/23/2023 02:30 PM |  |  | SAH8200 | Portable Ultrasound System | 18-143 | Scanning Systems, Ultrasonic, Portable | WK3D03 | 340440 | SonoSite div of FUJIFILM | P08778 |  | CATH LAB | 1105-30610 | SAH-Cardiac Cath Lab | 04 | Inventory | SAH | SAH
1810520 | 09/20/2023 04:50 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Spoke with Trinity Biomedical-Device needs a new board | Crystal Schmitt | CRSCH | EMPLOYEE | 09/25/2023 04:50 PM | 10/25/2023 02:15 PM |  |  | SAH4194 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248862 | 312187 | Fresenius Medical Care North A | 2008 T |  | Dialysis T15 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1818557 | 10/03/2023 11:17 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Leaking | Crystal Schmitt | CRSCH | EMPLOYEE | 10/08/2023 11:17 AM | 10/17/2023 12:50 PM |  |  | SAH11396 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1397868 | 199732 | Mar Cor Purification A Cantel | WRO300H |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1819212 | 10/06/2023 08:41 AM | CM | Corrective Maintenance | IP | In Progress | Device rear housing damaged. Found during EOC rounds. | Crystal Schmitt | CRSCH | EMPLOYEE | 10/11/2023 08:41 AM | 10/17/2023 10:37 AM | INA | Not Repaired-Device Disposed | SAH3137 | Fetal and Vascular Doppler | 10-429 | Detectors, Blood Flow, Ultrasonic | EAL1929 | 458815 | Natus Neurology Inc | 100 |  | 3E Oncology. | 1105-12245 | SAH-Oncology Unit AT6 | 04 | Inventory | SAH | SAH
1819636 | 10/09/2023 11:24 AM | CM | Corrective Maintenance | IP | In Progress | Shop-Stock order October 2023 | Crystal Schmitt | CRSCH | EMPLOYEE | 10/14/2023 11:24 AM | 10/20/2023 10:35 AM |  |  | SHOP-STOCK | Shop Stock | CC-056 | Environment | Shop Stock | 000 | Misc-Numbers | Stock |  | SAH Hospital | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1821139 | 10/11/2023 02:43 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1118Loc: Dialysis Stock Room Machine stopped due to Alarm. Pneumatic Alarm - M30.1 / Balloon Valve Leak | Crystal Schmitt | CRSCH | EMPLOYEE | 10/16/2023 02:43 AM | 10/30/2023 03:10 PM |  |  | L42237 | Fluid Exchanger/Heater | 11-226 | Peritoneal Dialysis Units | LC101013 | 312187 | Fresenius Medical Care North A | 180343 |  | Dialysis Stock Room | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1821490 | 10/12/2023 12:06 PM | CM | Corrective Maintenance | IP | In Progress | GEHC came onsite to rearrange CMU. Documenting all corrections that needed to be made after CMU was re-arranged. | Crystal Schmitt | CRSCH | EMPLOYEE | 10/17/2023 12:06 PM | 10/18/2023 07:34 AM |  |  | SPEC-PRO | Department Environment Rounds | CC-056 | Environment | Special Project | 000 | Misc-Numbers | ENVIRONMENT |  | SAH - Clinical Engineering | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1822860 | 10/20/2023 07:44 AM | CM | Corrective Maintenance | OF | Out For Repair | 727 825 1021Broken transport monitor - fell off a shelf and is cracked@ front desk | Crystal Schmitt | CRSCH | EMPLOYEE | 10/25/2023 07:44 AM | 10/25/2023 09:58 AM |  |  | SAH12088 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT21242066SA | 440964 | GE Healthcare Technologies | B450 |  | TRANSPORT | 1105-20100 | SAH-Anesthesiology | IU | Active | SAH | SAH
1823628 | 10/24/2023 01:50 PM | CM | Corrective Maintenance | OF | Out For Repair | Reading incorrect ground resistance | Crystal Schmitt | CRSCH | EMPLOYEE | 10/29/2023 01:50 PM | 10/26/2023 09:20 AM |  |  | SAH8877 | Electrical Safety Analyzer | CC-040 | Test Equipment | 5942040 | 178848 | Fluke Biomedical | ESA609 |  |  | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1824313 | 10/27/2023 12:25 PM | CM | Corrective Maintenance | DU | Device Unavailable | x171141/Heart monitor touchscreen is not functioning. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/01/2023 12:25 PM | 11/02/2023 09:35 AM |  |  | SAH11105 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT20482166SA | 440964 | GE Healthcare Technologies | b450 V3 |  | E304 | 1105-12154 | SAH-Cardiovascular Tele PT3 | IU | Active | SAH | SAH
1828936 | 11/01/2023 01:10 PM | CM | Corrective Maintenance | IP | In Progress | Special Projects and assignments November 2023 | Crystal Schmitt | CRSCH | EMPLOYEE | 11/06/2023 12:10 PM | 11/02/2023 03:50 PM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1828948 | 11/01/2023 01:50 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1287. The machine is making noise from compressor. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/06/2023 12:50 PM | 11/02/2023 04:09 PM |  |  | SAH3887 | High-level Scope Disinfector | 11-279 | Disinfectors, Liquid Germicide, Flexible Endoscope | 2633608 | 276750 | Olympus America Inc | OER-PRO |  | Endo Scope Room | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1828950 | 11/01/2023 01:51 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1287. Fluid sensor needs to be replaced. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/06/2023 12:51 PM | 11/02/2023 04:16 PM |  |  | SAH3886 | High-level Scope Disinfector | 11-279 | Disinfectors, Liquid Germicide, Flexible Endoscope | 2633605 | 276750 | Olympus America Inc | OER-PRO |  | Endo Decon | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1829136 | 11/02/2023 12:10 PM | CM | Corrective Maintenance | IP | In Progress | PM failed due to physical damage, needs new front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/07/2023 11:10 AM | 11/02/2023 12:34 PM |  |  | SAH3229 | LP20e Defibrillator/Pacer | 17-882B | Defibrillator/Pacemakers, External, Biphasic | 48234721 | 452575 | Physio-Control Inc | LP20E |  | SAH-3rd EP lab | 1105-13100 | SAH-Intensive Care Unit | LS | High Risk | SAH | SAH
1829724 | 11/04/2023 02:03 PM | CM | Corrective Maintenance | OP | Open | 727-825-1118 / Blood pump making noise; blood found in the pump; V105V108 error. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/09/2023 01:03 PM | 11/04/2023 02:03 PM |  |  | SAH4494 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 1T0S256665 | 312187 | Fresenius Medical Care North A | 2008 T |  | Hemodialysis Lab | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH