## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1857801 | 01/11/2024 09:12 AM | CM | Corrective Maintenance | IP | In Progress | RPT project OR | Crystal Schmitt | CRSCH | EMPLOYEE | 01/16/2024 09:12 AM | 03/27/2024 10:50 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1872943 | 02/20/2024 05:34 PM | CM | Corrective Maintenance | WP | Awaiting Parts | This device and other system components are plugged into a power transformer, which is plugged into a non-approved power strip. CM to correct power strip issue. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 05:34 PM | 06/24/2024 12:00 PM |  |  | SAH4002 | Stimulus Generator Unit | 17-898 | Monitoring Systems, Physiologic, Cardiac Electrophysiology | 4417-SGU/SC397 | 427969 | Bard Peripheral Vascular Div C | MP3008 |  | EP LAB | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1885408 | 03/22/2024 12:25 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Found during PM-Bottom corner LCD is dark and corner of front chassis damaged. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 12:25 PM | 06/17/2024 09:52 AM |  |  | SAH5687 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311035HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 09 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1885828 | 03/25/2024 03:56 PM | CM | Corrective Maintenance | IP | In Progress | 7278251287CO2 cylinder tank holder install needed, on siteEndo storage room 2 | Crystal Schmitt | CRSCH | EMPLOYEE | 03/30/2024 03:56 PM | 04/09/2024 02:30 PM |  |  | SAH5531 | Endo Motorized Travel Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 2000RFWDII082217A | 453029 | PHS West Inc | Ergo Express |  | na | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1892294 | 04/04/2024 06:26 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM Failed-USB port damaged-unable to block out port or connect any USB. Bottom corner of display is dark-display needs to be replaced. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/09/2024 06:26 PM | 06/17/2024 09:15 AM |  |  | SAH8034 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417321139HA | 440964 | GE Healthcare Technologies | B650 |  | EPACU01 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1892711 | 04/08/2024 11:42 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Found during PM-Damaged front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 04/13/2024 11:42 AM | 06/24/2024 01:58 PM |  |  | SAH5608 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311028HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 16 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1892850 | 04/09/2024 08:39 AM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Front chassis is damaged-clip attachment is broken off. 6/17-Found rear housing also damaged., cracked in multiple places. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/14/2024 08:39 AM | 06/20/2024 10:10 AM |  |  | SAH4953 | Bispectral Index BIS Module | 20-783 | Physiologic Monitor Modules, Electroencephalography | SGS19130002HA | 440964 | GE Healthcare Technologies | E-BIS-01 | CVICU | cvICU | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1906860 | 05/16/2024 08:56 AM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Set point temp out of range. Only reaches 39.3C. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/21/2024 08:56 AM | 05/16/2024 11:01 AM |  |  | L42239 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 91996 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1911478 | 05/17/2024 11:43 AM | CM | Corrective Maintenance | IP | In Progress | Found during PM-Caster breaks not functional. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/22/2024 11:43 AM | 05/17/2024 12:42 PM |  |  | SAH6364 | Mobile Exam Light | 33-257 | Lights, Examination, Stand-Alone, High End | 0819637 | 107071 | Medical Illumination Internat | 061523 |  |  | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1912625 | 05/25/2024 10:24 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Device reading "low voltage". Power management board issue | Crystal Schmitt | CRSCH | EMPLOYEE | 05/30/2024 10:24 AM | 05/28/2024 10:30 AM |  |  | SAH11072 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT20492117SA | 440964 | GE Healthcare Technologies | b450 V3 |  | EX06 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1913056 | 05/28/2024 11:25 AM | EN | Easy Net | IP | In Progress | when scanning patient it states there is no order for patient Note: This device is being relocated from 6N to Cardiac Medical II-3E once the repairs are completed. JM | Crystal Schmitt | CRSCH | EMPLOYEE | 06/02/2024 11:25 AM | 05/28/2024 06:59 PM |  |  | SAH1885 | EKG Mulitchannel Interp | 16-231 | Electrocardiographs, Multichannel, Interpretive | SKJ17210068PA | 440964 | GE Healthcare Technologies | MAC 5500HD |  | 6 N / LOC 901 | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1913215 | 05/28/2024 03:59 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damaged housing | Crystal Schmitt | CRSCH | EMPLOYEE | 06/02/2024 03:59 PM | 06/05/2024 09:55 AM |  |  | L49545 | Laparoscopic Insufflator | 16-849 | Insufflators, Laparoscopic | C03601 | 457057 | Medivators Corp A Cantel Medic | EGA-501 | ENDO | Endo 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1919344 | 06/05/2024 11:28 AM | CM | Corrective Maintenance | IP | In Progress | Special Projects and assignments June 2024. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/10/2024 11:28 AM | 06/25/2024 10:00 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1920404 | 06/06/2024 01:49 PM | CM | Corrective Maintenance | OF | Out For Repair | 7278251287 - not reading correctly | Crystal Schmitt | CRSCH | EMPLOYEE | 06/11/2024 01:49 PM | 06/25/2024 06:00 PM |  |  | SAH9200 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A1241956 | 107875 | Exergen Corp | TAT-5000 |  | ENDOSCOPY | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1920437 | 06/06/2024 02:52 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Found during Steris PM-Plastic cover needs to be replaced and both breaks need to be fixed | Crystal Schmitt | CRSCH | EMPLOYEE | 06/11/2024 02:52 PM | 06/17/2024 07:58 AM |  |  | SAH5336 | Surgical Boom | 22-613 | Facility Booms, Ceiling-Mounted | 0420017002 | 107114 | STERIS Corp | HarmonyAIR |  | Endo 4 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1920440 | 06/06/2024 02:54 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Found during Steris PM-Plastic cover needs replaced and breaks not working | Crystal Schmitt | CRSCH | EMPLOYEE | 06/11/2024 02:54 PM | 06/28/2024 04:25 PM |  |  | SAH5339 | Surgical Boom | 22-613 | Facility Booms, Ceiling-Mounted | 0420017009 | 107114 | STERIS Corp | HarmonyAIR |  | Endo 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1920447 | 06/06/2024 02:56 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Found during Steris PM-Plastic cover needs replaced and breaks not working | Crystal Schmitt | CRSCH | EMPLOYEE | 06/11/2024 02:56 PM | 06/06/2024 03:10 PM |  |  | SAH5338 | Surgical Boom | 22-613 | Facility Booms, Ceiling-Mounted | 0420017007 | 107114 | STERIS Corp | HarmonyAIR |  | Endo 3 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1920452 | 06/06/2024 03:02 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Found during Steris PM-Plastic cover needs replaced | Crystal Schmitt | CRSCH | EMPLOYEE | 06/11/2024 03:02 PM | 06/07/2024 02:37 PM |  |  | SAH5342 | Surgical Boom | 22-613 | Facility Booms, Ceiling-Mounted | 0420017008 | 107114 | STERIS Corp | HarmonyAIR |  | Endo Rm 3 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1920739 | 06/07/2024 12:48 PM | CM | Corrective Maintenance | AF | Awaiting FSR | Trinity Biomedical performing software update all Dialysis machines | Crystal Schmitt | CRSCH | EMPLOYEE | 06/12/2024 12:48 PM | 06/07/2024 12:10 PM |  |  | SAH6551 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 7T0S213143 | 312187 | Fresenius Medical Care North A | 2008 T |  | TRAVEL #2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1922533 | 06/12/2024 09:37 PM | EN | Easy Net | OF | Out For Repair | Numbers are not coming up or displaying correctly | Crystal Schmitt | CRSCH | EMPLOYEE | 06/17/2024 09:37 PM | 06/18/2024 12:55 PM |  |  | SAH5523 | Hemodynamic Fluid Monitor | 10-613 | Cardiac Output Units | V52016010528 | 454355 | Cheetah Medical | Starling SV | ICU | ICU | 1105-13100 | SAH-Intensive Care Unit | 04 | Inventory | SAH | SAH
1923025 | 06/15/2024 09:46 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118 / Failing temperature / Water room | Crystal Schmitt | CRSCH | EMPLOYEE | 06/20/2024 09:46 AM | 06/17/2024 08:21 AM |  |  | SAH4188 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248874 | 312187 | Fresenius Medical Care North A | 2008 T |  | T10 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1923348 | 06/18/2024 08:11 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1039. Micro stem does not have output. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/23/2024 08:11 AM | 06/20/2024 09:38 AM |  |  | SAH5276 | Nerve Block Stimulators | 16-252 | Stimulators, Electrical, Peripheral Nerve, Block Monitoring | NSNBC74339 | 105216 | Neuro Technology Inc | MICROSTIM |  | 3 NW | 1105-13100 | SAH-Intensive Care Unit | 04 | Inventory | SAH | SAH
1923436 | 06/18/2024 12:11 PM | CM | Corrective Maintenance | OF | Out For Repair | Device reading low temp, and missing key ring | Crystal Schmitt | CRSCH | EMPLOYEE | 06/23/2024 12:11 PM | 06/20/2024 08:41 AM |  |  | SAH9199 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A1241927 | 107875 | Exergen Corp | TAT-5000 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1923437 | 06/18/2024 12:12 PM | CM | Corrective Maintenance | OF | Out For Repair | Device reading low temp and missing key ring | Crystal Schmitt | CRSCH | EMPLOYEE | 06/23/2024 12:12 PM | 06/20/2024 08:46 AM |  |  | SAH9197 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A1241951 | 107875 | Exergen Corp | TAT-5000 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1923720 | 06/19/2024 06:19 PM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | Machine T17, SAH 4492 unable to test due to conductivity low, bicarb low error despite troubleshooting, wand care, and machine turn off/on. Machine also had similar issue last month. Was able to vinegar and heat machine prior to tagging/svc request. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/24/2024 06:19 PM | 06/20/2024 12:55 AM |  |  | SAH4492 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 1T0S256689 | 312187 | Fresenius Medical Care North A | 2008 T |  | Dialysis Water Room | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1923763 | 06/20/2024 08:31 AM | EN | Easy Net | OF | Out For Repair | Myron not calibrating. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/25/2024 08:31 AM | 06/26/2024 12:05 PM |  |  | SAH7017 | Dialysate Conductivity Meters | 11-206 | Electrical Conductivity Meters, Dialysate | 497166 | 101527 | Myron L Co | D-1 |  | HEMODIALYSIS LAB | 1105-32100 | SAH-Hemodialysis | 04 | Inventory | SAH | SAH
1923794 | 06/20/2024 10:05 AM | CM | Corrective Maintenance | IP | In Progress | Error USB device not recognized. Vendor ticket number is mst0632832. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/25/2024 10:05 AM | 06/26/2024 10:15 AM |  |  | SAH9608 | Data Recorder Kit | 34-626 | Recorders, Physiologic, pH, Upper Gastrointestinal Tract, Ingestible Capsule | DR30720763 | 391192 | Medtronic Given Imaging Inc | FGS-0347 |  | Endoscopy Phys Dictation | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1924425 | 06/24/2024 01:47 PM | CM | Corrective Maintenance | IP | In Progress | Leakage tester used with Tee clean needs to be sent for calibration and added to AIMs. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/29/2024 01:47 PM | 06/24/2024 09:30 AM |  |  | ENV-EKG | Department Environment Rounds | CC-056 | Environment | Environment | 000 | Misc-Numbers | ENVIRONMENT |  | SAH-EKG-Echo Dept | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1924745 | 06/25/2024 06:32 PM | CM | Corrective Maintenance | AD | Awaiting Delivery of Equipment | PM failed-Device will not turn on for vendor during PM | Crystal Schmitt | CRSCH | EMPLOYEE | 06/30/2024 06:32 PM | 06/28/2024 03:30 PM |  |  | SAH6767 | Front End Hemodynamic Monitor | 12-648 | Monitoring Systems, Physiologic, Cardiac Catheterization | US1810001524 | 452291 | Philips Interventional Patient | FC2010 |  | OR2 | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1925250 | 06/28/2024 12:33 PM | CM | Corrective Maintenance | IP | In Progress | Device alarming insufficient water two hours into treatment. | Crystal Schmitt | CRSCH | EMPLOYEE | 07/03/2024 12:33 PM | 06/28/2024 12:38 PM |  |  | SAH11396 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1397868 | 199732 | Mar Cor Purification A Cantel | WRO300H |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1925255 | 06/28/2024 12:54 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 7278251118- Giving error saying :Insufficient water during patient dialysis treatment | Crystal Schmitt | CRSCH | EMPLOYEE | 07/03/2024 12:54 PM | 06/28/2024 01:10 PM |  |  | SAH11396 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1397868 | 199732 | Mar Cor Purification A Cantel | WRO300H |  | Dialysis Water rm | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1925277 | 06/28/2024 02:42 PM | CM | Corrective Maintenance | IP | In Progress | PM failed-Damaged rear housing and front housing-needs to go in for depot repair. | Crystal Schmitt | CRSCH | EMPLOYEE | 07/03/2024 02:42 PM | 06/28/2024 04:10 PM |  |  | SAH11749 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNE19262042SA | 440964 | GE Healthcare Technologies | B450 |  | TRANSPORT | 1105-13102 | SAH-Intensive Care Unit II 3NW | IU | Active | SAH | SAH
1925300 | 06/28/2024 05:51 PM | CM | Corrective Maintenance | IP | In Progress | Documenting adding Aeroscout tags to portable monitors in Endo prep rooms. | Crystal Schmitt | CRSCH | EMPLOYEE | 07/03/2024 05:51 PM | 06/28/2024 06:06 PM |  |  | ENV-ENDO | Department Environment Rounds | CC-056 | Environment | Environment | 000 | Misc-Numbers | ENVIRONMENT |  | SAH-Endoscopy | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1925376 | 06/30/2024 08:45 AM | EN | Easy Net | OP | Open | The blood pressure cable is not working on the machine. | Crystal Schmitt | CRSCH | EMPLOYEE | 07/05/2024 08:45 AM | 06/30/2024 08:45 AM |  |  | SAH5693 | Hemodynamic Fluid Monitor | 10-613 | Cardiac Output Units | V5201510142 | 454355 | Cheetah Medical | Starling SV | ICU | ICU | 1105-13100 | SAH-Intensive Care Unit | 04 | Inventory | SAH | SAH
1925394 | 06/30/2024 03:48 PM | EN | Easy Net | OP | Open | EKG leads are reading blank and unable to record any activity | Crystal Schmitt | CRSCH | EMPLOYEE | 07/05/2024 03:48 PM | 06/30/2024 03:48 PM |  |  | SAH9275 | EKG Mulitchannel Interp | 16-231 | Electrocardiographs, Multichannel, Interpretive | VU223200128SA | 440964 | GE Healthcare Technologies | MAC VU360 | ICU | ICU | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH