## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1829136 | 11/02/2023 12:10 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed due to physical damage, needs new front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/07/2023 11:10 AM | 01/12/2024 10:26 AM |  |  | SAH3229 | LP20e Defibrillator/Pacer | 17-882B | Defibrillator/Pacemakers, External, Biphasic | 48234721 | 452575 | Physio-Control Inc | LP20E |  | SAH-3rd EP lab | 1105-13100 | SAH-Intensive Care Unit | LS | High Risk | SAH | SAH
1835648 | 11/14/2023 12:53 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/19/2023 12:53 PM | 11/15/2023 09:36 AM |  |  | L41843 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84662 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1836456 | 11/17/2023 06:33 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 7278251118 - error high/low flow, temp & negative pressure | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 06:33 AM | 01/23/2024 10:05 AM | COM | Component Failure-NOT Battery | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water RM AT 2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1836574 | 11/17/2023 01:13 PM | CM | Corrective Maintenance | OF | Out For Repair | Device error indicating board failure. Rep on-site. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 01:13 PM | 11/29/2023 09:01 AM |  |  | SAH9152 | Electrosurgical Unit | 18-232 | Electrosurgical Units, Monopolar/Bipolar, Argon-Enhanced Coagulation | 11535108 | 108422 | ERBE USA Inc | 10160-100 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1836624 | 11/17/2023 03:20 PM | CM | Corrective Maintenance | OF | Out For Repair | Error message. Erbe tech onsite troubleshooting. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 03:20 PM | 11/29/2023 09:06 AM |  |  | SAH9153 | Argon Plasma Coag Module | 17-738 | Gas Delivery Units, Argon-Enhanced Coagulation | 11532372 | 108422 | ERBE USA Inc | APC3 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1838000 | 11/28/2023 02:05 PM | EN | Easy Net | WP | Awaiting Parts | Manoscan machine in ENDO preop room 5, has been failing to pressurize during calibration of probe. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/03/2023 02:05 PM | 12/29/2023 09:57 AM |  |  | SAH5651 | ManoScan Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 6071 | 391192 | Medtronic Given Imaging Inc | 8039 |  | Endoscopy, preop room 5 | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1845993 | 12/08/2023 04:30 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Needs new level sensor | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 04:30 PM | 02/01/2024 10:43 AM |  |  | SAH6111 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYARY026 | 451492 | Medivance Inc A Bard Medical C | 5000 |  | ER- | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1846703 | 12/13/2023 08:37 AM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | Signature suite; cannot select different scenery. Reset the signature suite generator in the Steris Rack without resolution. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/18/2023 08:37 AM | 01/25/2024 09:55 AM |  |  | SAH11475 | OR Video Integration Systems | 20-388 | Automation Systems, Operating Room | R221790458 | 107114 | STERIS Corp | HexaVue System |  | OR 17 | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1846871 | 12/13/2023 03:21 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Device turned off and will not power back on | Crystal Schmitt | CRSCH | EMPLOYEE | 12/18/2023 03:21 PM | 01/30/2024 12:40 PM |  |  | SAH8863 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1407677 | 199732 | Mar Cor Purification A Cantel | MILLENIUM HX |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1853577 | 01/02/2024 08:26 AM | CM | Corrective Maintenance | IP | In Progress | Special Projects and Assignments January 2024 | Crystal Schmitt | CRSCH | EMPLOYEE | 01/07/2024 08:26 AM | 01/29/2024 04:50 PM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1856674 | 01/04/2024 04:28 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM Failed-Component loose inside of device | Crystal Schmitt | CRSCH | EMPLOYEE | 01/09/2024 04:28 PM | 01/05/2024 09:10 AM |  |  | SAH8533 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20190431SA | 440964 | GE Healthcare Technologies | 2016793-003 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1857424 | 01/09/2024 11:06 AM | CM | Corrective Maintenance | OF | Out For Repair | Broken screen | Crystal Schmitt | CRSCH | EMPLOYEE | 01/14/2024 11:06 AM | 01/16/2024 08:07 AM |  |  | SAH9318 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A1386483 | 107875 | Exergen Corp | TAT-5000 |  | ENDOSCOPY | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1857801 | 01/11/2024 09:12 AM | CM | Corrective Maintenance | IP | In Progress | RPT project OR | Crystal Schmitt | CRSCH | EMPLOYEE | 01/16/2024 09:12 AM | 02/01/2024 11:00 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1857999 | 01/11/2024 05:12 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | T-13 - RO machine 202 conductivity alarm - 727-825-1118 (RO9) | Crystal Schmitt | CRSCH | EMPLOYEE | 01/16/2024 05:12 PM | 02/01/2024 10:50 AM |  |  | SAH4489 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1397865 | 199732 | Mar Cor Purification A Cantel | WRO300H |  | R9 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1859725 | 01/19/2024 07:30 PM | CM | Corrective Maintenance | IP | In Progress | Device missing thermometer. | Crystal Schmitt | CRSCH | EMPLOYEE | 01/24/2024 07:30 PM | 01/19/2024 07:32 PM | ACC | Accessory Failure | SAH2893 | Electronic Thermometer | 14-032 | Thermometers, Electronic | 19500366 | 101850 | Welch Allyn Inc A Subsidiary | 692 |  |  | 1105-23100 | SAH-Emergency Service | 04 | Inventory | SAH | SAH
1859745 | 01/20/2024 08:11 AM | CM | Corrective Maintenance | OF | Out For Repair | 727-825-1706 Displaying: Monitor will not start Contact GE SupportRoom 25 | Crystal Schmitt | CRSCH | EMPLOYEE | 01/25/2024 08:11 AM | 01/25/2024 04:14 PM |  |  | SAH11237 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT20492125SA | 440964 | GE Healthcare Technologies | b450 V3 |  | EX25 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1859993 | 01/22/2024 10:12 AM | CM | Corrective Maintenance | OF | Out For Repair | Device error "service gas module". | Crystal Schmitt | CRSCH | EMPLOYEE | 01/27/2024 10:12 AM | 01/25/2024 04:16 PM |  |  | SAH6195 | Airway Gas Module w/Agent Id | 20-778 | Physiologic Monitor Modules, Anesthetic Gas | SGV19190279HA | 440964 | GE Healthcare Technologies | E-sCAiO-00 |  |  | 1105-20100 | SAH-Anesthesiology | IU | Active | SAH | SAH
1860155 | 01/22/2024 05:42 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-NIBP leak | Crystal Schmitt | CRSCH | EMPLOYEE | 01/27/2024 05:42 PM | 02/01/2024 02:36 PM |  |  | SAH11457 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX22330555SA | 440964 | GE Healthcare Technologies | 2016793-003 |  | AT2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1860437 | 01/23/2024 03:22 PM | CM | Corrective Maintenance | OF | Out For Repair | Error code 41 low flow. Artic Sun failing functional check after returning from PM. | Crystal Schmitt | CRSCH | EMPLOYEE | 01/28/2024 03:22 PM | 01/26/2024 11:58 AM |  |  | SAH5374 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYCSY555 | 451492 | Medivance Inc A Bard Medical C | 5000 | ICU | ICU | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1860729 | 01/25/2024 07:41 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-656-5267. Not able to scroll between screens. | Crystal Schmitt | CRSCH | EMPLOYEE | 01/30/2024 07:41 AM | 01/25/2024 09:50 AM |  |  | SAH6358 | OR Video Integration Systems | 20-388 | Automation Systems, Operating Room | R0952 | 107114 | STERIS Corp | CXPS System |  | ENDO 4 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1861936 | 01/26/2024 05:43 PM | CM | Corrective Maintenance | IP | In Progress | PM failed-AC inlet damaged. Loaner being returned without exchange. | Crystal Schmitt | CRSCH | EMPLOYEE | 01/31/2024 05:43 PM | 01/26/2024 05:50 PM |  |  | L42190 | Fluid Exchanger/Heater | 11-226 | Peritoneal Dialysis Units | LC110943 | 312187 | Fresenius Medical Care North A | 180343 |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1868523 | 02/01/2024 02:39 PM | CM | Corrective Maintenance | IP | In Progress | EOC rounds Cath lab | Crystal Schmitt | CRSCH | EMPLOYEE | 02/06/2024 02:39 PM | 02/01/2024 02:40 PM |  |  | ENV-CATHLB | Department Environment Rounds | CC-056 | Environment | Environment | 000 | Misc-Numbers | ENVIRONMENT |  | SAH-Cardiac Cath Lab | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1868525 | 02/01/2024 02:48 PM | CM | Corrective Maintenance | IP | In Progress | NIBP connector damaged | Crystal Schmitt | CRSCH | EMPLOYEE | 02/06/2024 02:48 PM | 02/01/2024 12:30 PM |  |  | SAH11399 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SA317380054SA | 440964 | GE Healthcare Technologies | 2016793-003 |  | EC 15 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1868562 | 02/01/2024 04:30 PM | EN | Easy Net | OP | Open | Doppler does not turn on | Crystal Schmitt | CRSCH | EMPLOYEE | 02/06/2024 04:30 PM | 02/01/2024 04:30 PM |  |  | SAH6362 | Blood Flow Detector | 10-429 | Detectors, Blood Flow, Ultrasonic | NSNBC19256 | 102998 | Parks Medical Electronics Inc | 811B |  | PACU 2 | 1105-21500 | SAH-PACU | 04 | Inventory | SAH | SAH