## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1835648 | 11/14/2023 12:53 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/19/2023 12:53 PM | 02/26/2024 11:35 AM |  |  | L41843 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84662 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1836574 | 11/17/2023 01:13 PM | CM | Corrective Maintenance | IP | In Progress | Device error indicating board failure. Rep on-site. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 01:13 PM | 02/20/2024 05:00 PM |  |  | SAH9473 | Electrosurgical Unit | 18-232 | Electrosurgical Units, Monopolar/Bipolar, Argon-Enhanced Coagulation | 11535108 | 108422 | ERBE USA Inc | 10160-100 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1836624 | 11/17/2023 03:20 PM | CM | Corrective Maintenance | IP | In Progress | Error message. Erbe tech onsite troubleshooting. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 03:20 PM | 02/20/2024 05:00 PM | 210 | Could not duplicate | SAH9153 | Argon Plasma Coag Module | 17-738 | Gas Delivery Units, Argon-Enhanced Coagulation | 11532372 | 108422 | ERBE USA Inc | APC3 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1838000 | 11/28/2023 02:05 PM | EN | Easy Net | WP | Awaiting Parts | Manoscan machine in ENDO preop room 5, has been failing to pressurize during calibration of probe. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/03/2023 02:05 PM | 12/29/2023 09:57 AM |  |  | SAH5651 | ManoScan Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 6071 | 391192 | Medtronic Given Imaging Inc | 8039 |  | Endoscopy, preop room 5 | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1845993 | 12/08/2023 04:30 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Needs new level sensor | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 04:30 PM | 01/26/2024 11:58 AM |  |  | SAH6111 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYARY026 | 451492 | Medivance Inc A Bard Medical C | 5000 |  | ER- | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1846871 | 12/13/2023 03:21 PM | CM | Corrective Maintenance | AF | Awaiting FSR | Device turned off and will not power back on | Crystal Schmitt | CRSCH | EMPLOYEE | 12/18/2023 03:21 PM | 02/27/2024 04:40 PM |  |  | SAH8863 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1407677 | 199732 | Mar Cor Purification A Cantel | MILLENIUM HX |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1856674 | 01/04/2024 04:28 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM Failed-Component loose inside of device | Crystal Schmitt | CRSCH | EMPLOYEE | 01/09/2024 04:28 PM | 02/12/2024 07:15 AM |  |  | SAH8533 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20190431SA | 440964 | GE Healthcare Technologies | 2016793-003 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1857801 | 01/11/2024 09:12 AM | CM | Corrective Maintenance | IP | In Progress | RPT project OR | Crystal Schmitt | CRSCH | EMPLOYEE | 01/16/2024 09:12 AM | 02/22/2024 11:30 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1860437 | 01/23/2024 03:22 PM | CM | Corrective Maintenance | OF | Out For Repair | Error code 41 low flow. Artic Sun failing functional check after returning from PM. | Crystal Schmitt | CRSCH | EMPLOYEE | 01/28/2024 03:22 PM | 01/26/2024 11:58 AM |  |  | SAH5374 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYCSY555 | 451492 | Medivance Inc A Bard Medical C | 5000 | ICU | ICU | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1862213 | 01/29/2024 03:49 PM | CM | Corrective Maintenance | WP | Awaiting Parts | 727.502.4391 / PH Meter receives error during calibration process / Chemistry Dept, 1st Floor | Crystal Schmitt | CRSCH | EMPLOYEE | 02/03/2024 03:49 PM | 02/22/2024 03:22 PM |  |  | SAH11375 | Ph Meters | 15-164 | pH Meters | AE95014245 | 375242 | Fisher Healthcare Div Thermo S | AE 150 |  | SAH-Chemistry | 1105-31000 | SAH-Laboratory Services | 04 | Inventory | SAH | SAH
1869284 | 02/06/2024 05:42 PM | CM | Corrective Maintenance | IP | In Progress | Special Projects and Assignments February 2024 | Crystal Schmitt | CRSCH | EMPLOYEE | 02/11/2024 05:42 PM | 02/23/2024 03:20 PM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1869476 | 02/07/2024 11:17 AM | CM | Corrective Maintenance | AF | Awaiting FSR | Software version does not match other telemetry servers. Issue with power distribution to servers and receivers in MDF closet. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/12/2024 11:17 AM | 02/20/2024 11:05 AM |  |  | SAH3336 | Telemetry Server | 21-973 | Network Servers | SEE09100133GA | 440964 | GE Healthcare Technologies | MP100R |  | MDF-AT2 | 1105-18415 | SAH-Neuroscience Unit 5N | IU | Active | SAH | SAH
1870133 | 02/10/2024 05:48 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Two keypads are not workingLOC: Dialysis Water Room | Crystal Schmitt | CRSCH | EMPLOYEE | 02/15/2024 05:48 PM | 02/22/2024 09:15 AM |  |  | SAH1864 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 8T0S234256 | 312187 | Fresenius Medical Care North A | 2008 T |  | Dialysis Water Room | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1870657 | 02/13/2024 03:33 PM | CM | Corrective Maintenance | OF | Out For Repair | PM Failed-Electrical safety failed | Crystal Schmitt | CRSCH | EMPLOYEE | 02/18/2024 03:33 PM | 02/26/2024 09:39 AM |  |  | L42225 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84661 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1870663 | 02/13/2024 03:44 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Damaged top case-Found during PM | Crystal Schmitt | CRSCH | EMPLOYEE | 02/18/2024 03:44 PM | 02/13/2024 03:48 PM |  |  | SAH11033 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20490799SA | 440964 | GE Healthcare Technologies | 2016793-003 |  | ICU BED 11 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1870687 | 02/13/2024 05:22 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Top case needs to be replaced | Crystal Schmitt | CRSCH | EMPLOYEE | 02/18/2024 05:22 PM | 02/15/2024 07:32 AM |  |  | SAH11064 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20490754SA | 440964 | GE Healthcare Technologies | 2016793-003 |  | CVICU9 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1871138 | 02/15/2024 04:29 PM | CM | Corrective Maintenance | IP | In Progress | PB980 training | Crystal Schmitt | CRSCH | EMPLOYEE | 02/20/2024 04:29 PM | 02/15/2024 05:25 PM |  |  | TRAINING | Used for Training Work Orders | CC-056 | Environment | NSN | 000 | Misc-Numbers | Training |  |  | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1871351 | 02/16/2024 11:11 AM | CM | Corrective Maintenance | IP | In Progress | Device being disposed; Dialysis switching to Baxter PD machines | Crystal Schmitt | CRSCH | EMPLOYEE | 02/21/2024 11:11 AM | 02/22/2024 01:05 PM |  |  | L42237 | Fluid Exchanger/Heater | 11-226 | Peritoneal Dialysis Units | LC101013 | 312187 | Fresenius Medical Care North A | 180343 |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1871352 | 02/16/2024 11:13 AM | CM | Corrective Maintenance | IP | In Progress | Device being disposed; Dialysis switching to Baxter PD machines | Crystal Schmitt | CRSCH | EMPLOYEE | 02/21/2024 11:13 AM | 02/22/2024 12:55 PM | INA | Not Repaired-Device Disposed | L42023 | Fluid Exchanger/Heater | 11-226 | Peritoneal Dialysis Units | LC122577 | 312187 | Fresenius Medical Care North A | 180343 |  | Storage Rm | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1872905 | 02/20/2024 02:28 PM | CM | Corrective Maintenance | WP | Awaiting Parts | 727-825-1287 / Low pressure on the pump. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 02:28 PM | 02/23/2024 07:52 AM |  |  | SAH5288 | Endoscope Flushing Pump | 13-192 | Pumps | 21735979 | 276750 | Olympus America Inc | OFP-2 |  | Endo Procedure Rm 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1872943 | 02/20/2024 05:34 PM | CM | Corrective Maintenance | WP | Awaiting Parts | This device and other system components are plugged into a power transformer, which is plugged into a non-approved power strip. CM to correct power strip issue. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 05:34 PM | 02/20/2024 05:38 PM |  |  | SAH4002 | Stimulus Generator Unit | 17-898 | Monitoring Systems, Physiologic, Cardiac Electrophysiology | 4417-SGU/SC397 | 427969 | Bard Peripheral Vascular Div C | MP3008 |  | EP LAB | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1874209 | 02/26/2024 11:12 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118, Not passing negative pressure and tempLOC: Water Room, Dialysis | Crystal Schmitt | CRSCH | EMPLOYEE | 03/02/2024 11:12 AM | 02/27/2024 01:36 PM |  |  | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water Room, Dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1874232 | 02/26/2024 12:10 PM | CM | Corrective Maintenance | IP | In Progress | PM failed-Damaged housing; both top and bottom plastic cracked. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/02/2024 12:10 PM | 02/26/2024 11:15 AM |  |  | L41902 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 87724 | 418683 | Arizant Healthcare Inc | 245 |  | SAH-Surgery | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1874913 | 02/27/2024 12:50 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed- "Replace water trap" message did not appear and "Sample gas out" error appears when port not blocked. Pump is loud and turns off intermittently. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/03/2024 12:50 PM | 02/27/2024 01:01 PM |  |  | SAH2699 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8530508 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | PACU 15 | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH