## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1845993 | 12/08/2023 04:30 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Needs new level sensor3/4/24-Failed "inspect fluid delivery line" check for one valve | Crystal Schmitt | CRSCH | EMPLOYEE | 12/13/2023 04:30 PM | 04/16/2024 01:55 PM |  |  | SAH6111 | Temperature Management System | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | DYARY026 | 451492 | Medivance Inc A Bard Medical C | 5000 |  | ER- | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1857801 | 01/11/2024 09:12 AM | CM | Corrective Maintenance | IP | In Progress | RPT project OR | Crystal Schmitt | CRSCH | EMPLOYEE | 01/16/2024 09:12 AM | 03/27/2024 10:50 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1862213 | 01/29/2024 03:49 PM | CM | Corrective Maintenance | WP | Awaiting Parts | 727.502.4391 / PH Meter receives error during calibration process / Chemistry Dept, 1st Floor | Crystal Schmitt | CRSCH | EMPLOYEE | 02/03/2024 03:49 PM | 02/22/2024 03:22 PM |  |  | SAH11375 | Ph Meters | 15-164 | pH Meters | AE95014245 | 375242 | Fisher Healthcare Div Thermo S | AE 150 |  | SAH-Chemistry | 1105-31000 | SAH-Laboratory Services | 04 | Inventory | SAH | SAH
1870687 | 02/13/2024 05:22 PM | CM | Corrective Maintenance | IP | In Progress | Top case needs to be replaced | Crystal Schmitt | CRSCH | EMPLOYEE | 02/18/2024 05:22 PM | 04/01/2024 07:29 AM | PHY | Physical Damage | SAH11064 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20490754SA | 440964 | GE Healthcare Technologies | 2016793-003 |  | CVICU9 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1871138 | 02/15/2024 04:29 PM | CM | Corrective Maintenance | IP | In Progress | PB980 training | Crystal Schmitt | CRSCH | EMPLOYEE | 02/20/2024 04:29 PM | 02/15/2024 05:25 PM |  |  | TRAINING | Used for Training Work Orders | CC-056 | Environment | NSN | 000 | Misc-Numbers | Training |  |  | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1872943 | 02/20/2024 05:34 PM | CM | Corrective Maintenance | WP | Awaiting Parts | This device and other system components are plugged into a power transformer, which is plugged into a non-approved power strip. CM to correct power strip issue. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 05:34 PM | 02/20/2024 05:38 PM |  |  | SAH4002 | Stimulus Generator Unit | 17-898 | Monitoring Systems, Physiologic, Cardiac Electrophysiology | 4417-SGU/SC397 | 427969 | Bard Peripheral Vascular Div C | MP3008 |  | EP LAB | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1874209 | 02/26/2024 11:12 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118, Not passing negative pressure and tempLOC: Water Room, Dialysis | Crystal Schmitt | CRSCH | EMPLOYEE | 03/02/2024 11:12 AM | 02/27/2024 01:36 PM |  |  | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water Room, Dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1882410 | 03/12/2024 10:55 AM | CM | Corrective Maintenance | IP | In Progress | Need key tether installed | Crystal Schmitt | CRSCH | EMPLOYEE | 03/17/2024 10:55 AM | 03/12/2024 10:59 AM |  |  | SAH5531 | Endo Motorized Travel Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 2000RFWDII082217A | 453029 | PHS West Inc | Ergo Express |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1883259 | 03/18/2024 03:10 AM | EN | Easy Net | AF | Awaiting FSR | T7RO with "insufficient water alarms" and "internal service alarms" Bay 3 HD suite portable hooked up to wall box at the time. No water issues when setting up the machine (rinsing, priming, testing etc.) but as soon as treatment started the RO wasn"t getting enough water. Machine tagged out in water room and Paul with biomed made aware. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/23/2024 03:10 AM | 04/03/2024 07:52 AM |  |  | SAH11396 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1397868 | 199732 | Mar Cor Purification A Cantel | WRO300H |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1883634 | 03/19/2024 08:03 PM | CM | Corrective Maintenance | IP | In Progress | Found power cord/charger damaged during department PMs | Crystal Schmitt | CRSCH | EMPLOYEE | 03/24/2024 08:03 PM | 03/19/2024 08:05 PM |  |  | SAH3370 | Printer | 22-203 | Printers, Computer/Network, Ink-Jet | U64418-G6W123430 | 473839 | Brother International | MW-260MFi |  | PIC TEAM | 1105-37200 | SAH-IV Therapy | 04 | Inventory | SAH | SAH
1885219 | 03/21/2024 12:44 PM | CM | Corrective Maintenance | IP | In Progress | PM failed-Failed pulse width, pulse frequency and stimulus duration test. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/26/2024 12:44 PM | 04/18/2024 07:50 AM |  |  | SAH5393 | Electroshock Unit | 11-484 | Stimulators, Electrical, Brain, Convulsive Therapy | 42685 | 104776 | Somatics LLC | System IV |  | SAH-PACU | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH
1885387 | 03/22/2024 11:01 AM | CM | Corrective Maintenance | AD | Awaiting Delivery of Equipment | 727-825-1125 or 727-692-9317/Not running blood pressure. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 11:01 AM | 04/09/2024 04:45 PM |  |  | SAH5881 | Front End Hemodynamic Monitor | 12-648 | Monitoring Systems, Physiologic, Cardiac Catheterization | US1550001478 | 452291 | Philips Interventional Patient | FC2020 |  | Cath Lab 2 | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1885408 | 03/22/2024 12:25 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Found during PM-Bottom corner LCD is dark and corner of front chassis damaged. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 12:25 PM | 03/22/2024 12:25 PM |  |  | SAH5687 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311035HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 09 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1885424 | 03/22/2024 01:17 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to top and bottom housing | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 01:17 PM | 03/22/2024 01:01 PM |  |  | L42214 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 82323 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1885828 | 03/25/2024 03:56 PM | CM | Corrective Maintenance | IP | In Progress | 7278251287CO2 cylinder tank holder install needed, on siteEndo storage room 2 | Crystal Schmitt | CRSCH | EMPLOYEE | 03/30/2024 03:56 PM | 04/09/2024 02:30 PM |  |  | SAH5531 | Endo Motorized Travel Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 2000RFWDII082217A | 453029 | PHS West Inc | Ergo Express |  | na | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1886842 | 03/26/2024 11:54 AM | CM | Corrective Maintenance | IP | In Progress | Time did not update on large wall display | Crystal Schmitt | CRSCH | EMPLOYEE | 03/31/2024 11:54 AM | 04/15/2024 01:00 PM |  |  | SAH9311 | OR Video Integration Systems | 20-388 | Automation Systems, Operating Room | R221790457 | 107114 | STERIS Corp | HexaVue System |  | Surgery Core#3 near OR15 | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1888539 | 03/29/2024 01:44 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727 825 1118RO Machine: Service Required. Component failure | Crystal Schmitt | CRSCH | EMPLOYEE | 04/03/2024 01:44 PM | 04/12/2024 10:50 AM |  |  | SAH8862 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1407678 | 199732 | Mar Cor Purification A Cantel | MILLENIUM HX |  | SAH / Dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1889662 | 04/01/2024 10:20 AM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Housing damaged | Crystal Schmitt | CRSCH | EMPLOYEE | 04/06/2024 10:20 AM | 04/16/2024 11:15 AM |  |  | SAH9589 | Laparoscopic Insufflator | 16-849 | Insufflators, Laparoscopic | C03602 | 457057 | Medivators Corp A Cantel Medic | EGA-501 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1891643 | 04/02/2024 02:03 PM | CM | Corrective Maintenance | IP | In Progress | 7278251287- SAH Dictation room - Device giving error saying "DR3 battery needs to be replaced" | Crystal Schmitt | CRSCH | EMPLOYEE | 04/07/2024 02:03 PM | 04/09/2024 02:15 PM |  |  | SAH12058 | Data Recorder Kit | 34-626 | Recorders, Physiologic, pH, Upper Gastrointestinal Tract, Ingestible Capsule | DR30823485 | 391192 | Medtronic Given Imaging Inc | FGS-0347 |  | SAH-Endo/Dictation Room | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1891847 | 04/03/2024 07:45 AM | CM | Corrective Maintenance | IP | In Progress | Special Projects and Assignments April 2024 | Crystal Schmitt | CRSCH | EMPLOYEE | 04/08/2024 07:45 AM | 04/23/2024 11:20 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1892294 | 04/04/2024 06:26 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM Failed-USB port damaged-unable to block out port or connect any USB. Bottom corner of display is dark-display needs to be replaced. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/09/2024 06:26 PM | 04/09/2024 09:10 AM |  |  | SAH8034 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417321139HA | 440964 | GE Healthcare Technologies | B650 |  | EPACU01 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1892326 | 04/05/2024 06:50 AM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | V105 V108 STUCK CLOSED. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/10/2024 06:50 AM | 04/05/2024 07:59 AM |  |  | SAH4194 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248862 | 312187 | Fresenius Medical Care North A | 2008 T |  | water room. | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1892477 | 04/05/2024 05:50 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Leak test failed | Crystal Schmitt | CRSCH | EMPLOYEE | 04/10/2024 05:50 PM | 04/23/2024 12:20 PM |  |  | SAH4475 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8806475 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | ICU7 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1892711 | 04/08/2024 11:42 AM | CM | Corrective Maintenance | IP | In Progress | Found during PM-Damaged front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 04/13/2024 11:42 AM | 04/09/2024 07:55 AM |  |  | SAH5608 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311028HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 16 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1892850 | 04/09/2024 08:39 AM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Front chassis is damaged-clip attachment is broken off. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/14/2024 08:39 AM | 04/09/2024 08:20 AM |  |  | SAH4953 | Bispectral Index BIS Module | 20-783 | Physiologic Monitor Modules, Electroencephalography | SGS19130002HA | 440964 | GE Healthcare Technologies | E-BIS-01 | CVICU | cvICU | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1892971 | 04/09/2024 11:02 AM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Needs front housing replaced due to damaged plastic. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/14/2024 11:02 AM | 04/23/2024 11:27 AM |  |  | SAH11863 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNE17360054HA | 440964 | GE Healthcare Technologies | B450 |  | ICU storage shelf | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1893772 | 04/12/2024 06:01 AM | CM | Corrective Maintenance | AC | Acknowledged | 727-825-1118 \ Turned off with ''no water'' error \ Loc: dialysis water room | Crystal Schmitt | CRSCH | EMPLOYEE | 04/17/2024 06:01 AM | 04/15/2024 09:45 AM |  |  | SAH1865 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 9T0S236481 | 312187 | Fresenius Medical Care North A | 2008 T |  | T7 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1893898 | 04/12/2024 11:07 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Trinity Biomedical informed that device cart is leaking. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/17/2024 11:07 AM | 04/16/2024 02:23 PM |  |  | SAH2545 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1377811 | 199732 | Mar Cor Purification A Cantel | WRO300H |  | RO Machine #8 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1894348 | 04/16/2024 07:10 AM | CM | Corrective Maintenance | OF | Out For Repair | Problem: piece broken off of back of deviceLOC: EndoPhone:727.825.1287 | Crystal Schmitt | CRSCH | EMPLOYEE | 04/21/2024 07:10 AM | 04/22/2024 08:35 AM |  |  | SAH4453 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A973741 | 107875 | Exergen Corp | TAT-5000 | ENDO | Endo 2 | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1894486 | 04/16/2024 10:58 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Device has a cracked rear bezel and an issue with the display backlight. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/21/2024 10:58 AM | 04/22/2024 09:41 AM |  |  | SAH7374 | Non-Invasive Ventilator | 20-746 | Ventilators, Noninvasive Positive Pressure | 100040164 | 454161 | Philips Respironics | V60 |  | AH_V60_06 | 1105-35000 | SAH-Respiratory Care Services 2nd | LS | High Risk | SAH | SAH
1894517 | 04/16/2024 12:31 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Pressure gague broken, weak suction | Crystal Schmitt | CRSCH | EMPLOYEE | 04/21/2024 12:31 PM | 04/16/2024 12:10 PM |  |  | SAH5686 | Portable Suction | 27-657 | Aspirators, Airway | PD505288 | 453450 | DeVilbiss Healthcare Inc | 7305P-D |  | ICU in front of room 13 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1897585 | 04/19/2024 02:09 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Endo leak tester has a leak in the tubing | Crystal Schmitt | CRSCH | EMPLOYEE | 04/24/2024 02:09 PM | 04/22/2024 07:58 AM |  |  | SAH5701 | Endoscope- Pneumatic Tester | 23-981 | Testers, Endoscope, Leak/Fluid Detection | 7659553 | 276750 | Olympus America Inc | MU-1 | ENDO | Endo 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1899275 | 04/22/2024 09:50 AM | CM | Corrective Maintenance | AF | Awaiting FSR | 317-695-1633/Printer (SAH3276) not found error is being displayed. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/24/2024 09:50 AM | 04/24/2024 11:06 AM |  |  | SAH5232 | OR Video Integration Systems | 20-388 | Automation Systems, Operating Room | R1140 | 107114 | STERIS Corp | CXPS System |  | OR Pump room for OR 4&6 | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1899447 | 04/23/2024 08:41 AM | CM | Corrective Maintenance | IP | In Progress | 727 452 2477Tee Probe cleaner. Message: Doors not closed but doors are closed. Unable to use machine | Crystal Schmitt | CRSCH | EMPLOYEE | 04/25/2024 08:41 AM | 04/23/2024 11:00 AM |  |  | SAH11284 | TEE Probe Disinfector | 26-654 | Disinfectors, Liquid Germicide, Transesophageal Echocardiography Probe | 20100164 | 456998 | CS Medical LLC | TCD100S8 |  | Echo Dept | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH