## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1857801 | 01/11/2024 09:12 AM | CM | Corrective Maintenance | IP | In Progress | RPT project OR | Crystal Schmitt | CRSCH | EMPLOYEE | 01/16/2024 09:12 AM | 03/27/2024 10:50 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1871138 | 02/15/2024 04:29 PM | CM | Corrective Maintenance | IP | In Progress | PB980 training | Crystal Schmitt | CRSCH | EMPLOYEE | 02/20/2024 04:29 PM | 05/20/2024 09:25 AM |  |  | TRAINING | Used for Training Work Orders | CC-056 | Environment | NSN | 000 | Misc-Numbers | Training |  |  | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1872943 | 02/20/2024 05:34 PM | CM | Corrective Maintenance | WP | Awaiting Parts | This device and other system components are plugged into a power transformer, which is plugged into a non-approved power strip. CM to correct power strip issue. | Crystal Schmitt | CRSCH | EMPLOYEE | 02/25/2024 05:34 PM | 02/20/2024 05:38 PM |  |  | SAH4002 | Stimulus Generator Unit | 17-898 | Monitoring Systems, Physiologic, Cardiac Electrophysiology | 4417-SGU/SC397 | 427969 | Bard Peripheral Vascular Div C | MP3008 |  | EP LAB | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1885219 | 03/21/2024 12:44 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Failed pulse width, pulse frequency and stimulus duration test. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/26/2024 12:44 PM | 05/16/2024 02:25 PM |  |  | SAH5393 | Electroshock Unit | 11-484 | Stimulators, Electrical, Brain, Convulsive Therapy | 42685 | 104776 | Somatics LLC | System IV |  | SAH-PACU | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH
1885408 | 03/22/2024 12:25 PM | CM | Corrective Maintenance | WP | Awaiting Parts | Found during PM-Bottom corner LCD is dark and corner of front chassis damaged. | Crystal Schmitt | CRSCH | EMPLOYEE | 03/27/2024 12:25 PM | 03/22/2024 12:25 PM |  |  | SAH5687 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311035HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 09 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1885828 | 03/25/2024 03:56 PM | CM | Corrective Maintenance | IP | In Progress | 7278251287CO2 cylinder tank holder install needed, on siteEndo storage room 2 | Crystal Schmitt | CRSCH | EMPLOYEE | 03/30/2024 03:56 PM | 04/09/2024 02:30 PM |  |  | SAH5531 | Endo Motorized Travel Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 2000RFWDII082217A | 453029 | PHS West Inc | Ergo Express |  | na | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1888539 | 03/29/2024 01:44 PM | CM | Corrective Maintenance | AF | Awaiting FSR | 727 825 1118RO Machine: Service Required. Component failure | Crystal Schmitt | CRSCH | EMPLOYEE | 04/03/2024 01:44 PM | 05/30/2024 07:58 AM |  |  | SAH8862 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1407678 | 199732 | Mar Cor Purification A Cantel | MILLENIUM HX |  | SAH / Dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1892294 | 04/04/2024 06:26 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM Failed-USB port damaged-unable to block out port or connect any USB. Bottom corner of display is dark-display needs to be replaced. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/09/2024 06:26 PM | 04/09/2024 09:10 AM |  |  | SAH8034 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417321139HA | 440964 | GE Healthcare Technologies | B650 |  | EPACU01 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1892711 | 04/08/2024 11:42 AM | CM | Corrective Maintenance | IP | In Progress | Found during PM-Damaged front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 04/13/2024 11:42 AM | 04/09/2024 07:55 AM |  |  | SAH5608 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SK417311028HA | 440964 | GE Healthcare Technologies | B650 |  | ICU BED 16 | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1892850 | 04/09/2024 08:39 AM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Front chassis is damaged-clip attachment is broken off. | Crystal Schmitt | CRSCH | EMPLOYEE | 04/14/2024 08:39 AM | 05/30/2024 06:19 PM |  |  | SAH4953 | Bispectral Index BIS Module | 20-783 | Physiologic Monitor Modules, Electroencephalography | SGS19130002HA | 440964 | GE Healthcare Technologies | E-BIS-01 | CVICU | cvICU | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1905270 | 05/07/2024 05:14 PM | CM | Corrective Maintenance | OF | Out For Repair | Monitor giving error message "Monitor will not start. Contact GE Service for support". | Crystal Schmitt | CRSCH | EMPLOYEE | 05/12/2024 05:14 PM | 05/30/2024 05:30 PM |  |  | SAH9713 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT20492125SA | 440964 | GE Healthcare Technologies | b450 V3 |  | EX25 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1906860 | 05/16/2024 08:56 AM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Set point temp out of range. Only reaches 39.3C. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/21/2024 08:56 AM | 05/16/2024 11:01 AM |  |  | L42239 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 91996 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1911478 | 05/17/2024 11:43 AM | CM | Corrective Maintenance | IP | In Progress | Found during PM-Caster breaks not functional. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/22/2024 11:43 AM | 05/17/2024 12:42 PM |  |  | SAH6364 | Mobile Exam Light | 33-257 | Lights, Examination, Stand-Alone, High End | 0819637 | 107071 | Medical Illumination Internat | 061523 |  |  | 1105-30610 | SAH-Cardiac Cath Lab | IU | Active | SAH | SAH
1911662 | 05/20/2024 08:27 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Found crack in front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 05/25/2024 08:27 AM | 05/23/2024 11:50 AM |  |  | SAH8862 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1407678 | 199732 | Mar Cor Purification A Cantel | MILLENIUM HX |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1911984 | 05/21/2024 11:21 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118 Machine (T-6) was reviewed for safety but requesting ticket to biomed. During CRRT and patient expired while on machine 5/16. ¦ SAH Hemodialysis Water Room | Crystal Schmitt | CRSCH | EMPLOYEE | 05/26/2024 11:21 AM | 06/05/2024 08:25 AM |  |  | SAH1864 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 8T0S234256 | 312187 | Fresenius Medical Care North A | 2008 T |  | T-6 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1912037 | 05/21/2024 01:52 PM | CM | Corrective Maintenance | IP | In Progress | 727-825-1287 / Pump not pumping properly, pumping intermittently. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/26/2024 01:52 PM | 05/21/2024 06:24 PM |  |  | SAH5756 | Endoscope Flushing Pump | 13-192 | Pumps | 21219041 | 276750 | Olympus America Inc | OFP |  | Endoscopy procedure rm 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1912595 | 05/24/2024 05:15 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damaged housing-bottom housing cracked | Crystal Schmitt | CRSCH | EMPLOYEE | 05/29/2024 05:15 PM | 05/30/2024 02:15 PM |  |  | L42274 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 104395 | 418683 | Arizant Healthcare Inc | 245 |  | SAH-Surgery | 1105-13102 | SAH-Intensive Care Unit II 3NW | IU | Active | SAH | SAH
1912625 | 05/25/2024 10:24 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Device reading "low voltage". Power management board issue | Crystal Schmitt | CRSCH | EMPLOYEE | 05/30/2024 10:24 AM | 05/28/2024 10:30 AM |  |  | SAH11072 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT20492117SA | 440964 | GE Healthcare Technologies | b450 V3 |  | EX06 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1913050 | 05/28/2024 11:20 AM | CM | Corrective Maintenance | IP | In Progress | 727.825.1030 - device keeps failing, vendor is aware, repaired it last week. When trying to clean a probe, it is failing. Manual cleans are not very healthful. | Crystal Schmitt | CRSCH | EMPLOYEE | 05/30/2024 11:20 AM | 05/31/2024 02:15 PM |  |  | SAH11284 | TEE Probe Disinfector | 26-654 | Disinfectors, Liquid Germicide, Transesophageal Echocardiography Probe | 20100164 | 456998 | CS Medical LLC | TCD100S8 |  | 3rd Floor - ECHO Dept. | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1913056 | 05/28/2024 11:25 AM | EN | Easy Net | IP | In Progress | when scanning patient it states there is no order for patient | Crystal Schmitt | CRSCH | EMPLOYEE | 06/02/2024 11:25 AM | 05/28/2024 06:59 PM |  |  | SAH1885 | EKG Mulitchannel Interp | 16-231 | Electrocardiographs, Multichannel, Interpretive | SKJ17210068PA | 440964 | GE Healthcare Technologies | MAC 5500HD |  | 6 N / LOC 901 | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1913215 | 05/28/2024 03:59 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damaged housing | Crystal Schmitt | CRSCH | EMPLOYEE | 06/02/2024 03:59 PM | 06/05/2024 09:55 AM |  |  | L49545 | Laparoscopic Insufflator | 16-849 | Insufflators, Laparoscopic | C03601 | 457057 | Medivators Corp A Cantel Medic | EGA-501 | ENDO | Endo 2 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1913725 | 05/31/2024 10:50 AM | CM | Corrective Maintenance | AF | Awaiting FSR | 727-825-1118. Safety check on machine. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/05/2024 10:50 AM | 06/05/2024 08:05 AM |  |  | SAH4191 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248652 | 312187 | Fresenius Medical Care North A | 2008 T |  | Dialysis Water Room | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1913736 | 05/31/2024 11:26 AM | CM | Corrective Maintenance | OF | Out For Repair | Initial inspection failed-Device not heating up to temp. Only reaches 37.1C. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/05/2024 11:26 AM | 06/05/2024 09:10 AM |  |  | L42406 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 076507 | 418683 | Arizant Healthcare Inc | 245 |  | Surgery | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1914463 | 05/31/2024 06:29 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed-Device batteries not holding a charge. Needs new batteries | Crystal Schmitt | CRSCH | EMPLOYEE | 06/05/2024 06:29 PM | 06/05/2024 09:17 AM |  |  | SAH6334 | Surgical Imaging Table | 16-600 | Tables, Imaging, Radiographic/Fluoroscopic | ISR3065 | 170423 | Image Diagnostics Inc | ISR G2 |  | COV Bed 3 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1916443 | 06/02/2024 10:29 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1118 / Feed line connection broke. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/07/2024 10:29 AM | 06/03/2024 10:34 AM |  |  | SAH4191 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248652 | 312187 | Fresenius Medical Care North A | 2008 T |  | AT2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1918322 | 06/03/2024 09:05 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1030TEE Clean machine: failing cleaning processSAH > Echo Dept > 3rd FL | Crystal Schmitt | CRSCH | EMPLOYEE | 06/05/2024 09:05 AM | 06/03/2024 03:22 PM |  |  | SAH11284 | TEE Probe Disinfector | 26-654 | Disinfectors, Liquid Germicide, Transesophageal Echocardiography Probe | 20100164 | 456998 | CS Medical LLC | TCD100S8 |  | SAH | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1919161 | 06/04/2024 11:40 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1039 / TM reported that the monitor suddenly started sending incorrect heart rate readings at 10:15am, but is still sending correct blood pressures. TM advised that the incorrect reading seem to be from the patient''s old room. / ICU 2, Bed 5 / | Crystal Schmitt | CRSCH | EMPLOYEE | 06/09/2024 11:40 AM | 06/05/2024 04:50 PM |  |  | SAH6248 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SQC19252125SA | 440964 | GE Healthcare Technologies | B650 |  | CVICU_05 | 1105-13102 | SAH-Intensive Care Unit II 3NW | IU | Active | SAH | SAH
1919344 | 06/05/2024 11:28 AM | CM | Corrective Maintenance | IP | In Progress | Special Projects and assignments June 2024. | Crystal Schmitt | CRSCH | EMPLOYEE | 06/10/2024 11:28 AM | 06/05/2024 11:31 AM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH