## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1809239 | 09/18/2023 08:04 AM | CM | Corrective Maintenance | IP | In Progress | 727 502 4411 US machine, Sonisite - Error Code: P...... | Crystal Schmitt | CRSCH | EMPLOYEE | 09/23/2023 08:04 AM | 10/23/2023 02:30 PM |  |  | SAH8200 | Portable Ultrasound System | 18-143 | Scanning Systems, Ultrasonic, Portable | WK3D03 | 340440 | SonoSite div of FUJIFILM | P08778 |  | CATH LAB | 1105-30610 | SAH-Cardiac Cath Lab | 04 | Inventory | SAH | SAH
1810520 | 09/20/2023 04:50 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Spoke with Trinity Biomedical-Device needs a new board | Crystal Schmitt | CRSCH | EMPLOYEE | 09/25/2023 04:50 PM | 10/25/2023 02:15 PM |  |  | SAH4194 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248862 | 312187 | Fresenius Medical Care North A | 2008 T |  | Dialysis T15 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1818073 | 10/02/2023 05:06 PM | CM | Corrective Maintenance | IP | In Progress | Special Projects and Assignments October 2023 | Crystal Schmitt | CRSCH | EMPLOYEE | 10/07/2023 05:06 PM | 10/25/2023 01:25 PM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1818557 | 10/03/2023 11:17 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Leaking | Crystal Schmitt | CRSCH | EMPLOYEE | 10/08/2023 11:17 AM | 10/17/2023 12:50 PM |  |  | SAH11396 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1397868 | 199732 | Mar Cor Purification A Cantel | WRO300H |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1819212 | 10/06/2023 08:41 AM | CM | Corrective Maintenance | IP | In Progress | Device rear housing damaged. Found during EOC rounds. | Crystal Schmitt | CRSCH | EMPLOYEE | 10/11/2023 08:41 AM | 10/17/2023 10:37 AM | INA | Not Repaired-Device Disposed | SAH3137 | Fetal and Vascular Doppler | 10-429 | Detectors, Blood Flow, Ultrasonic | EAL1929 | 458815 | Natus Neurology Inc | 100 |  | 3E Oncology. | 1105-12245 | SAH-Oncology Unit AT6 | 04 | Inventory | SAH | SAH
1819636 | 10/09/2023 11:24 AM | CM | Corrective Maintenance | IP | In Progress | Shop-Stock order October 2023 | Crystal Schmitt | CRSCH | EMPLOYEE | 10/14/2023 11:24 AM | 10/20/2023 10:35 AM |  |  | SHOP-STOCK | Shop Stock | CC-056 | Environment | Shop Stock | 000 | Misc-Numbers | Stock |  | SAH Hospital | 1105-88570 | SAH-Clinical Engineering | 04 | Inventory | SAH | SAH
1821139 | 10/11/2023 02:43 AM | CM | Corrective Maintenance | IP | In Progress | 727-825-1118Loc: Dialysis Stock Room Machine stopped due to Alarm. Pneumatic Alarm - M30.1 / Balloon Valve Leak | Crystal Schmitt | CRSCH | EMPLOYEE | 10/16/2023 02:43 AM | 10/13/2023 09:40 AM |  |  | L42237 | Fluid Exchanger/Heater | 11-226 | Peritoneal Dialysis Units | LC101013 | 312187 | Fresenius Medical Care North A | 180343 |  | Dialysis Stock Room | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1821490 | 10/12/2023 12:06 PM | CM | Corrective Maintenance | IP | In Progress | GEHC came onsite to rearrange CMU. Documenting all corrections that needed to be made after CMU was re-arranged. | Crystal Schmitt | CRSCH | EMPLOYEE | 10/17/2023 12:06 PM | 10/18/2023 07:34 AM |  |  | SPEC-PRO | Department Environment Rounds | CC-056 | Environment | Special Project | 000 | Misc-Numbers | ENVIRONMENT |  | SAH - Clinical Engineering | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1822860 | 10/20/2023 07:44 AM | CM | Corrective Maintenance | OF | Out For Repair | 727 825 1021Broken transport monitor - fell off a shelf and is cracked@ front desk | Crystal Schmitt | CRSCH | EMPLOYEE | 10/25/2023 07:44 AM | 10/25/2023 09:58 AM |  |  | SAH12088 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT21242066SA | 440964 | GE Healthcare Technologies | B450 |  | TRANSPORT | 1105-20100 | SAH-Anesthesiology | IU | Active | SAH | SAH
1823628 | 10/24/2023 01:50 PM | CM | Corrective Maintenance | OF | Out For Repair | Reading incorrect ground resistance | Crystal Schmitt | CRSCH | EMPLOYEE | 10/29/2023 01:50 PM | 10/26/2023 09:20 AM |  |  | SAH8877 | Electrical Safety Analyzer | CC-040 | Test Equipment | 5942040 | 178848 | Fluke Biomedical | ESA609 |  |  | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1823865 | 10/25/2023 05:00 PM | CM | Corrective Maintenance | IP | In Progress | 727-825-1051 / Bladder scanner is not scanner is not scanning correctly. It is giving incorrect readings. / 5 South East, behind the nurse''s station / | Crystal Schmitt | CRSCH | EMPLOYEE | 10/30/2023 05:00 PM | 10/26/2023 08:20 AM |  |  | SAH6087 | Bladder Vol and Mass Ultrasnd | 34-698 | Scanning Systems, Ultrasonic, Small-Parts, Urinary Bladder | R52M40QNZAP | 494232 | Signostics Inc | USCAN |  | 5East ( PROBE SN RB02731 ) | 1105-12295 | SAH-Med/Surg Overflow 5E | IU | Active | SAH | SAH
1823908 | 10/26/2023 06:57 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118 / Flow Inlet error message/ Dialysis Unite Water room | Crystal Schmitt | CRSCH | EMPLOYEE | 10/31/2023 06:57 AM | 10/26/2023 08:50 AM |  |  | SAH4192 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248900 | 312187 | Fresenius Medical Care North A | 2008 T |  | T-13 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1824032 | 10/26/2023 01:49 PM | CM | Corrective Maintenance | IP | In Progress | blood leak error intermittent - V104/V108 error - AT 2 dialysis - 727-825-1118Paging CES again for issue | Crystal Schmitt | CRSCH | EMPLOYEE | 10/31/2023 01:49 PM | 10/26/2023 02:15 PM |  |  | SAH4187 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248902 | 312187 | Fresenius Medical Care North A | 2008 T |  | AT 2 dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH