## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1823628 | 10/24/2023 01:50 PM | CM | Corrective Maintenance | OF | Out For Repair | Reading incorrect ground resistance | Crystal Schmitt | CRSCH | EMPLOYEE | 10/29/2023 01:50 PM | 10/26/2023 09:20 AM |  |  | SAH8877 | Electrical Safety Analyzer | CC-040 | Test Equipment | 5942040 | 178848 | Fluke Biomedical | ESA609 |  |  | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1829136 | 11/02/2023 12:10 PM | CM | Corrective Maintenance | WP | Awaiting Parts | PM failed due to physical damage, needs new front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/07/2023 11:10 AM | 11/15/2023 12:20 PM |  |  | SAH3229 | LP20e Defibrillator/Pacer | 17-882B | Defibrillator/Pacemakers, External, Biphasic | 48234721 | 452575 | Physio-Control Inc | LP20E |  | SAH-3rd EP lab | 1105-13100 | SAH-Intensive Care Unit | LS | High Risk | SAH | SAH
1830745 | 11/10/2023 08:17 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 7278251287 - Unit not seeing 2 devices - 2nd floor 2 SE endo 5 | Crystal Schmitt | CRSCH | EMPLOYEE | 11/15/2023 08:17 AM | 11/10/2023 01:20 PM |  |  | SAH5343 | Intergration System | 20-388 | Automation Systems, Operating Room | RL0010F367C4EA | 107114 | STERIS Corp | IQ2800 |  | Under Desk Endo 5 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1835395 | 11/13/2023 10:22 AM | CM | Corrective Maintenance | OF | Out For Repair | Device says "Reconnect PDM" error message | Crystal Schmitt | CRSCH | EMPLOYEE | 11/18/2023 10:22 AM | 12/01/2023 10:00 AM |  |  | SAH11225 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20490696SA | 440964 | GE Healthcare Technologies | 2016793-003 |  | ICU9 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1835599 | 11/14/2023 10:08 AM | CM | Corrective Maintenance | AF | Awaiting FSR | 727-452-2477 / TEE Probe cleaner is not working - It is not doing what it was supposed to do to clean the probe. TM called the vendor, who advised to have CES call them at 919-255-9472. / Echo Dept. / | Crystal Schmitt | CRSCH | EMPLOYEE | 11/16/2023 10:08 AM | 11/29/2023 08:26 AM |  |  | SAH11284 | TEE Probe Disinfector | 26-654 | Disinfectors, Liquid Germicide, Transesophageal Echocardiography Probe | 20100164 | 456998 | CS Medical LLC | TCD100S8 |  | Echo Dept | 1105-30620 | SAH-Non Invasive Cardiology | IU | Active | SAH | SAH
1835648 | 11/14/2023 12:53 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Damage to housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/19/2023 12:53 PM | 11/15/2023 09:36 AM |  |  | L41843 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 84662 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1836373 | 11/16/2023 12:07 PM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Device not in range for set-point temperature | Crystal Schmitt | CRSCH | EMPLOYEE | 11/21/2023 12:07 PM | 11/17/2023 08:40 AM |  |  | L42276 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 076707 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1836442 | 11/16/2023 06:55 PM | CM | Corrective Maintenance | WP | Awaiting Parts | CO2 Monitor is making a rattling noise. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/21/2023 06:55 PM | 12/01/2023 12:50 PM |  |  | SAH2694 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8530515 | 440964 | GE Healthcare Technologies | E-MiniC-00 | PACU | PACU | 1105-21500 | SAH-PACU | IU | Active | SAH | SAH
1836456 | 11/17/2023 06:33 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 7278251118 - error high/low flow, temp & negative pressure | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 06:33 AM | 11/22/2023 08:55 AM |  |  | SAH4190 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248890 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water RM AT 2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1836574 | 11/17/2023 01:13 PM | CM | Corrective Maintenance | OF | Out For Repair | Device error indicating board failure. Rep on-site. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 01:13 PM | 11/29/2023 09:01 AM |  |  | SAH9152 | Electrosurgical Unit | 18-232 | Electrosurgical Units, Monopolar/Bipolar, Argon-Enhanced Coagulation | 11535108 | 108422 | ERBE USA Inc | 10160-100 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1836624 | 11/17/2023 03:20 PM | CM | Corrective Maintenance | OF | Out For Repair | Error message. Erbe tech onsite troubleshooting. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/22/2023 03:20 PM | 11/29/2023 09:06 AM |  |  | SAH9153 | Argon Plasma Coag Module | 17-738 | Gas Delivery Units, Argon-Enhanced Coagulation | 11532372 | 108422 | ERBE USA Inc | APC3 |  |  | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1836894 | 11/20/2023 09:37 AM | CM | Corrective Maintenance | OF | Out For Repair | PM failed-Not heating up to set point range. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/25/2023 09:37 AM | 11/22/2023 01:20 PM |  |  | L42271 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 106245 | 418683 | Arizant Healthcare Inc | 245 |  | Surgery | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1837670 | 11/27/2023 08:10 AM | CM | Corrective Maintenance | WP | Awaiting Parts | Not sending gas flow | Crystal Schmitt | CRSCH | EMPLOYEE | 12/02/2023 08:10 AM | 12/01/2023 12:37 PM |  |  | SAH11120 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8805151 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | EC 24 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1837723 | 11/27/2023 10:38 AM | CM | Corrective Maintenance | OF | Out For Repair | PM Failed-Device not heating up to temp. Out of range for set point temperature. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/02/2023 10:38 AM | 11/27/2023 11:07 AM |  |  | L42239 | Blood Warmer | 10-447 | Warming Units, Blood/Intravenous Solution | 91996 | 418683 | Arizant Healthcare Inc | 245 |  |  | 1105-21100 | SAH-Surgery | IU | Active | SAH | SAH
1837835 | 11/27/2023 05:20 PM | CM | Corrective Maintenance | IP | In Progress | PM failed-Leak test failed | Crystal Schmitt | CRSCH | EMPLOYEE | 12/02/2023 05:20 PM | 12/01/2023 01:30 PM |  |  | SAH6271 | Gas Module | 20-776 | Physiologic Monitor Modules, Exhaled Carbon Dioxide | 8529135 | 440964 | GE Healthcare Technologies | E-MiniC-00 |  | CVICU2 | 1105-13102 | SAH-Intensive Care Unit II 3NW | IU | Active | SAH | SAH
1838000 | 11/28/2023 02:05 PM | EN | Easy Net | WP | Awaiting Parts | Manoscan machine in ENDO preop room 5, has been failing to pressurize during calibration of probe. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/03/2023 02:05 PM | 11/29/2023 05:13 PM |  |  | SAH5651 | ManoScan Cart | 33-601 | Carts, Storage/Transport, Electromechanical/ Electronic Instrument/Equipment | 6071 | 391192 | Medtronic Given Imaging Inc | 8039 |  | Endoscopy, preop room 5 | 1105-21221 | SAH-Endoscopy (2SE) | 04 | Inventory | SAH | SAH
1840978 | 12/01/2023 07:52 PM | EN | Easy Net | OP | Open | only reads 85.5 for every patient...needs repair. Thanks! | Crystal Schmitt | CRSCH | EMPLOYEE | 12/06/2023 07:52 PM | 12/04/2023 03:52 PM |  |  | SAH4009 | Temporal Infrared Thermometer | 17-888 | Thermometers, Electronic, Infrared, Skin | A861278 | 107875 | Exergen Corp | TAT-5000 |  | SAH ICU-1 | 1105-13100 | SAH-Intensive Care Unit | 04 | Inventory | SAH | SAH
1841000 | 12/02/2023 11:16 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727 825 1118 ¦ Dialysis Machine Repeatedly Producing Alarm code: V105 ¦ SAH 2nd Floor Dialysis | Crystal Schmitt | CRSCH | EMPLOYEE | 12/07/2023 11:16 AM | 12/05/2023 08:23 AM |  |  | SAH4195 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248838 | 312187 | Fresenius Medical Care North A | 2008 T |  | SAH 2nd Floor Dialysis | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1841087 | 12/04/2023 12:53 AM | CM | Corrective Maintenance | IP | In Progress | Cooling machine is not reading probes, keeps saying check probe connection. LOC: ICU Nurses station 727-825-1020 | Crystal Schmitt | CRSCH | EMPLOYEE | 12/09/2023 12:53 AM | 12/05/2023 03:15 PM |  |  | SAH8921 | Hypo-Hyperthermia Unit | 12-074 | Warming/Cooling Units, Patient, Circulating-Fluid | 222-3-14796 | 513842 | Gentherm | 233 | ICU | ICU | 1105-13102 | SAH-Intensive Care Unit II 3NW | IU | Active | SAH | SAH
1841411 | 12/04/2023 09:35 AM | EN | Easy Net | OP | Open | Bladder scanner not sensing any urine in the bladder. Tried to scan with a different scanner that worked just fine. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/09/2023 09:35 AM | 12/04/2023 09:35 AM |  |  | SAH9191 | Bladder Vol and Mass Ultrasnd | 34-698 | Scanning Systems, Ultrasonic, Small-Parts, Urinary Bladder | C1520889 | 451928 | Verathon Inc | Bladderscan Prime PL |  | Electrical Closet | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1841466 | 12/04/2023 11:26 AM | EN | Easy Net | OP | Open | missing alarm lens on transport monitor | Crystal Schmitt | CRSCH | EMPLOYEE | 12/09/2023 11:26 AM | 12/04/2023 11:26 AM |  |  | SAH5349 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNE17360052HA | 440964 | GE Healthcare Technologies | B450 |  | ICU transport monitor | 1105-13100 | SAH-Intensive Care Unit | IU | Active | SAH | SAH
1841544 | 12/04/2023 02:00 PM | EN | Easy Net | OP | Open | Have two pieces of rental equipment coming in early tomorrow for 0730 cases -- an Ultrasound machine from Agiliti & an ESWL machine from Southern Litho for a 0730 case too. | Crystal Schmitt | CRSCH | EMPLOYEE | 12/09/2023 02:00 PM | 12/04/2023 02:00 PM |  |  | RENTAL | Rental | CC-056 | Environment | Rental | 000 | Misc-Numbers | RENTAL |  | Clinical Engineering | 1105-13100 | SAH-Intensive Care Unit | 04 | Inventory | SAH | SAH
1841837 | 12/05/2023 08:19 AM | EN | Easy Net | WP | Awaiting Parts | monitor not responding | Crystal Schmitt | CRSCH | EMPLOYEE | 12/10/2023 08:19 AM | 12/05/2023 09:25 AM |  |  | SAH11441 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT22270048HA | 440964 | GE Healthcare Technologies | B450 |  | AT2 T2 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1841866 | 12/05/2023 09:16 AM | CM | Corrective Maintenance | OP | Open | 7278251118 - V104/108 stuck closed | Crystal Schmitt | CRSCH | EMPLOYEE | 12/10/2023 09:16 AM | 12/05/2023 09:16 AM |  |  | SAH4188 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248874 | 312187 | Fresenius Medical Care North A | 2008 T |  | Water RM | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH