## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | DUE_DATETIME | STAT_DATETIME | FAILURE | FAILURE_DESC | TAG_NUMBER | DESCRIPTION | TYPE | TYPE_DESC | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | LOCATION_CODE | LOCATION_DESC | CHG_CTR | CHG_CTR_DESC | EQU_STATUS | EQU_STATUS_DESC | FACILITY | FACILITY_DESC
1809239 | 09/18/2023 08:04 AM | CM | Corrective Maintenance | IP | In Progress | 727 502 4411 US machine, Sonisite - Error Code: P...... | Crystal Schmitt | CRSCH | EMPLOYEE | 09/23/2023 08:04 AM | 11/13/2023 09:00 AM |  |  | SAH8200 | Portable Ultrasound System | 18-143 | Scanning Systems, Ultrasonic, Portable | WK3D03 | 340440 | SonoSite div of FUJIFILM | P08778 |  | CATH LAB | 1105-30610 | SAH-Cardiac Cath Lab | 04 | Inventory | SAH | SAH
1818557 | 10/03/2023 11:17 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | Leaking | Crystal Schmitt | CRSCH | EMPLOYEE | 10/08/2023 11:17 AM | 11/08/2023 07:46 AM |  |  | SAH11396 | Water Purification RO | 14-437 | Water Purification Systems, Reverse Osmosis | 1397868 | 199732 | Mar Cor Purification A Cantel | WRO300H |  |  | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1819212 | 10/06/2023 08:41 AM | CM | Corrective Maintenance | IP | In Progress | Device rear housing damaged. Found during EOC rounds. | Crystal Schmitt | CRSCH | EMPLOYEE | 10/11/2023 08:41 AM | 11/13/2023 12:20 PM | INA | Not Repaired-Device Disposed | SAH3137 | Fetal and Vascular Doppler | 10-429 | Detectors, Blood Flow, Ultrasonic | EAL1929 | 458815 | Natus Neurology Inc | 100 |  | 3E Oncology. | 1105-12245 | SAH-Oncology Unit AT6 | 04 | Inventory | SAH | SAH
1822860 | 10/20/2023 07:44 AM | CM | Corrective Maintenance | OF | Out For Repair | 727 825 1021Broken transport monitor - fell off a shelf and is cracked@ front desk | Crystal Schmitt | CRSCH | EMPLOYEE | 10/25/2023 07:44 AM | 10/25/2023 09:58 AM |  |  | SAH12088 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT21242066SA | 440964 | GE Healthcare Technologies | B450 |  | TRANSPORT | 1105-20100 | SAH-Anesthesiology | IU | Active | SAH | SAH
1823628 | 10/24/2023 01:50 PM | CM | Corrective Maintenance | OF | Out For Repair | Reading incorrect ground resistance | Crystal Schmitt | CRSCH | EMPLOYEE | 10/29/2023 01:50 PM | 10/26/2023 09:20 AM |  |  | SAH8877 | Electrical Safety Analyzer | CC-040 | Test Equipment | 5942040 | 178848 | Fluke Biomedical | ESA609 |  |  | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1824313 | 10/27/2023 12:25 PM | CM | Corrective Maintenance | DU | Device Unavailable | x171141/Heart monitor touchscreen is not functioning. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/01/2023 12:25 PM | 11/09/2023 02:20 PM |  |  | SAH11105 | Bedside Physiologic Monitor | 20-171 | Monitors, Physiologic, Multipurpose, Bedside, Modular | SNT20482166SA | 440964 | GE Healthcare Technologies | b450 V3 |  | E304 | 1105-12154 | SAH-Cardiovascular Tele PT3 | IU | Active | SAH | SAH
1828936 | 11/01/2023 01:10 PM | CM | Corrective Maintenance | IP | In Progress | Special Projects and assignments November 2023 | Crystal Schmitt | CRSCH | EMPLOYEE | 11/06/2023 12:10 PM | 11/10/2023 04:15 PM |  |  | SPECPRO-SAH | Special Projects | CC-056 | Environment | None | 000 | Misc-Numbers | SPECPRO |  | ***DO NOT USE***SPEC-PRO | 1105-88570 | SAH-Clinical Engineering | IU | Active | SAH | SAH
1829136 | 11/02/2023 12:10 PM | CM | Corrective Maintenance | IP | In Progress | PM failed due to physical damage, needs new front housing | Crystal Schmitt | CRSCH | EMPLOYEE | 11/07/2023 11:10 AM | 11/02/2023 12:34 PM |  |  | SAH3229 | LP20e Defibrillator/Pacer | 17-882B | Defibrillator/Pacemakers, External, Biphasic | 48234721 | 452575 | Physio-Control Inc | LP20E |  | SAH-3rd EP lab | 1105-13100 | SAH-Intensive Care Unit | LS | High Risk | SAH | SAH
1829724 | 11/04/2023 02:03 PM | CM | Corrective Maintenance | IP | In Progress | 727-825-1118 / Blood pump making noise; blood found in the pump; V105V108 error. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/09/2023 01:03 PM | 11/06/2023 07:27 AM |  |  | SAH4494 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 1T0S256665 | 312187 | Fresenius Medical Care North A | 2008 T |  | Hemodialysis Lab | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1830011 | 11/07/2023 06:01 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 727-825-1118Loc: DialysisUnit has Air Alarm, unbale to clear | Crystal Schmitt | CRSCH | EMPLOYEE | 11/12/2023 06:01 AM | 11/07/2023 07:47 AM |  |  | SAH4193 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248892 | 312187 | Fresenius Medical Care North A | 2008 T |  | T12 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1830745 | 11/10/2023 08:17 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 7278251287 - Unit not seeing 2 devices - 2nd floor 2 SE endo 5 | Crystal Schmitt | CRSCH | EMPLOYEE | 11/12/2023 08:17 AM | 11/10/2023 01:20 PM |  |  | SAH5343 | Intergration System | 20-388 | Automation Systems, Operating Room | RL0010F367C4EA | 107114 | STERIS Corp | IQ2800 |  | Under Desk Endo 5 | 1105-21221 | SAH-Endoscopy (2SE) | IU | Active | SAH | SAH
1830827 | 11/10/2023 01:59 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | 7278251118Leaking water from back panel and underneath machineDialysis water room | Crystal Schmitt | CRSCH | EMPLOYEE | 11/15/2023 01:59 PM | 11/13/2023 10:32 AM |  |  | SAH4192 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248900 | 312187 | Fresenius Medical Care North A | 2008 T |  | T-13 | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH
1830971 | 11/13/2023 07:05 AM | CM | Corrective Maintenance | IP | In Progress | QC fails during daily maintenance. Blood is showing negative. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/15/2023 07:05 AM | 11/13/2023 09:15 AM |  |  | SAH7324 | Semi-Auto Urine Analyzer | 16-886 | Analyzers, Laboratory, Urine | KPS25441707 | 451971 | Siemens Healthcare Diagnostics | Clinitek Advantus |  | Hemo | 1105-31000 | SAH-Laboratory Services | 04 | Inventory | SAH | SAH
1835395 | 11/13/2023 10:22 AM | CM | Corrective Maintenance | OF | Out For Repair | Device says "Reconnect PDM" error message | Crystal Schmitt | CRSCH | EMPLOYEE | 11/18/2023 10:22 AM | 11/13/2023 10:50 AM |  |  | SAH11225 | Patient Data Module | 20-784 | Physiologic Monitor Modules, Multiparameter | SPX20490696SA | 440964 | GE Healthcare Technologies | 2016793-003 |  | ICU9 | 1105-23100 | SAH-Emergency Service | IU | Active | SAH | SAH
1835412 | 11/13/2023 11:17 AM | CM | Corrective Maintenance | OP | Open | Flow inlet error. no flow error. high flow error. | Crystal Schmitt | CRSCH | EMPLOYEE | 11/18/2023 11:17 AM | 11/13/2023 11:17 AM |  |  | SAH4189 | Hemodialysis Unit | 11-218 | Hemodialysis Units | 0T0S248826 | 312187 | Fresenius Medical Care North A | 2008 T |  | 2nd Dialysis Water Room | 1105-32100 | SAH-Hemodialysis | IU | Active | SAH | SAH