## Work Orders
WO_NUMBER | REQST_DATETIME | WO_TYPE | WO_TYPE_DESC | WO_STATUS | WO_STATUS_DESC | REQUESTOR | CONTACT | WO_PROBLEM | TRADE_EMP_DESC | TRADE_EMP_CODE | TRADE_EMP_TYPE | PRIORITY | PRIORITY_DESC | TAG_NUMBER | DESCRIPTION | SERIAL_NUM | MANUFACTURER | MANUFACTURER_DESC | MODEL_NUM | FACILITY | FACILITY_DESC
1568048 | 02/09/2022 10:59 AM | SP | Special Projects | IW | Awaiting Onsite Vendor Maint | b119865 | Robert Nobrega / 7278594913 | Log4j Vulnerability Coordinate with vendor for remediation | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS6406 | Digital Portable X-ray System | DF2402000764WK | 439946 | GE Healthcare USA | 5555000-27 | SJS | SJS
1748847 | 04/19/2022 03:08 PM | SP | Special Projects | IW | Awaiting Onsite Vendor Maint | b119865 | ROBERT NOBREGA / 7278594913 | Remediate Axeda cybersecurity vulnerability following attached procedure | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC120108 | Computer Aided Detection Mammo | C5740 | 106487 | Hologic Inc | ASY-02354 | BOI | BOI
1772836 | 08/01/2023 07:46 AM | PM | Planned Maintenance | WP | Awaiting Parts |  |  | -Vendor PM - Next Scheduled Date: 02/01/2024 | KenMed Supply LLC | 569382 | VENDOR | 2 | PM Level 2 | WHH3279 | Surgical Table | 3500-3I-1286 | 102186 | Skytron Div The KMW Group | 3500 | WHH | WHH
1773865 | 06/08/2023 02:27 AM | EN | Easy Net | WP | Awaiting Parts | b144654 | 8632971805 | not connecting to program. | Med Stat Consulting SVCS | M10025 | VENDOR | 9 | 2 Routine | WHH7058 | Pump Control Unit 5.7 Screen | 15844815 | 378348 | BD | 8015LS | WHH | WHH
1777067 | 06/19/2023 09:59 AM | EN | Easy Net | IP | In Progress | b113787 | 863-297-1807 | IV pump wont scan | Med Stat Consulting SVCS | M10025 | VENDOR | 9 | 2 Routine | WHH2475 | Large Volume Pump Module | 4123275 | 378348 | BD | 8100 | WHH | WHH
1783113 | 07/06/2023 02:14 PM | CM | Corrective Maintenance | WP | Awaiting Parts | B105822 | Dianne Gee / 813-443-7163 | 8133028045 - Monitor on top broken off LOC: OR cove hallway | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4274 | Mobile C-Arm | E2-4440 | 453503 | GE Healthcare Surgery | 9900 ELITE | SJS | SJS
1805124 | 09/01/2023 08:15 AM | PM | Planned Maintenance | NP | PM Not Performed Missing |  |  | --- Ranger 245 PM Procedure --- - Next Scheduled Date: 12/01/2023 | KenMed Supply LLC | 569382 | VENDOR | 2 | PM Level 2 | WHH8867 | Blood Warmer | 3425 | 418683 | Arizant Healthcare Inc | 245 | WHH | WHH
1809151 | 09/16/2023 05:05 PM | CM | Corrective Maintenance | IP | In Progress | B125882 | Brian Bennett / 727-467-4700 | Joshua: 330-414-4561, Dominic Perna: 401-258-9606 / / Joshua Davis, service manager for Fujifilm, delivering a large part for MRI chiller tomorrow @ 8:30 AM to BOI Hampton Lakes. Asking if they can be let in on Sunday. Just need to unlock the door. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | NOTAG | Need Tag Number | NSNNOTAG | 000 | Misc-Numbers | No Tag | BMG | BMG
1885055 | 03/20/2024 01:45 PM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | EXD53563 | 813-356-7113 | field detector out of field/ center. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH28126 | Radiographic/Fluoroscopic Syst | 41C351685001 | 106973 | Shimadzu Medical Systems, Inc | SonialVision G4 | SJH | SJH
1891633 | 04/02/2024 01:46 PM | EN | Easy Net | IP | In Progress | b149445 | 8632971800 | Pump says channel error. Cant disconnect from pump | Med Stat Consulting SVCS | M10025 | VENDOR | 9 | 2 Routine | WHH2116 | Large Volume Pump Module | 4124694 | 378348 | BD | 8100 | WHH | WHH
1894758 | 04/17/2024 08:38 AM | CM | Corrective Maintenance | WP | Awaiting Parts | B117911 | Steven Kelly / | 813-302-8045/Will not move. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4274 | Mobile C-Arm | E2-4440 | 453503 | GE Healthcare Surgery | 9900 ELITE | SJS | SJS
1898825 | 05/01/2024 06:31 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Janet Pindral | 813.872.3949 | -Service Agreement PM - Next Scheduled Date: 11/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH4067 | 2D Digital Mammo | 81008110911 | 106487 | Hologic Inc | SDM-00001-2D | SJH | SJH
1899228 | 04/22/2024 07:39 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B157579 | Seth Pini / | 813-302-8743 / Repair request for lead shield mounted to ceiling / IR 1 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4415 | Angiographic/Interventional Sy | 121056 | 399199 | Siemens Medical Solutions USA | 10848282 | SJS | SJS
1899413 | 04/22/2024 10:28 PM | CM | Corrective Maintenance | CF | Conditionally Functional | JRB57478 | Jason Bowers / 813-870-4605 | Radiology 2nd floor : 813-870-4605: Plug is missing the ground | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH26543 | Digital Portable X-ray System | DF2401900172WK | 439946 | GE Healthcare USA | 5555000-47 | SJH | SJH
1901017 | 04/27/2024 01:13 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B096367 | Maria Gonzalez / 813-870-4605 | losing connection with board when placed behind patient 813-870-4000 x804605 OR x735494 LOC: Surgery dept 2nd floor equipment storage TM called to get status on WO and would like CES to reach out with a ETA on when they can expect to receive it to be fix. 8138704000 ext 803255. She requested to have WO changed to STAT. | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | SJH16157 | Digital Portable X-ray System | DF2401800163WK | 439946 | GE Healthcare USA | 5555000-27 | SJH | SJH
1901050 | 04/28/2024 01:41 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B077072 | Phil Lake / 813-870-4625 | 8138704625 / Need PO order for vendor repair / Catscan | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH8157 | CT Scanner | 8DA15X2004 | 101894 | Toshiba America Medical System | TSX-301C-6 | SJH | SJH
1901370 | 04/29/2024 03:25 PM | IN | Initial Inspection | IP | In Progress | b126836 | Jennah Judd / 727-315-5339 | REC 5/15 3D DIMENSION 5101-33510 BREAST CENTER 510122972 HOLOGIC QUOTE ATTACHED | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | NEW-EQUIP | New Equipment ADD | NSNNEW-EQUIPSJS | 000 | Misc-Numbers | New Equipment ADD | SJS | SJS
1903490 | 05/01/2024 11:10 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Val Relente | 813-302-8030 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 05/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJS4274 | Mobile C-Arm | E2-4440 | 453503 | GE Healthcare Surgery | 9900 ELITE | SJS | SJS
1904259 | 05/01/2024 12:06 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Meredith Alaynick | TBD | -Vendor PM - Next Scheduled Date: 05/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC171289 | MRI Water Chiller | W186509-001 | 553916 | Glen Dimplex Thermal Solutions | W02-10000-2P-NF-L-M | BOI | BOI
1904263 | 05/01/2024 12:06 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | n | n | -Service Agreement PM - Next Scheduled Date: 11/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC186779 | MRI 1.5 Tesla | 31351 | 399199 | Siemens Medical Solutions USA | 10018165 | BOI | BOI
1904491 | 05/02/2024 09:36 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B127848 | Amanda Logan / | 813.586.8508 / Mammo machine buttons need to be replaced. The machine is still usable in the meantime. / Bloomingdale HealthHub | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC146753 | 3D Digital Mammo | 3DM160100701 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1904634 | 04/29/2024 03:25 PM | IN | Initial Inspection | IW | Awaiting Onsite Vendor Maint | b126836 | Jennah Judd / 727-315-5339 | ETA 6/6/2024 FORESIGHT TIMS COMPUTER 5102-21221 SJS-ENDOSCOPY | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | NEW-EQUIP | New Equipment ADD | NSNNEW-EQUIPSJS | 000 | Misc-Numbers | New Equipment ADD | SJS | SJS
1906023 | 05/13/2024 08:34 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B119731 | Dannielle Alfonso / | 813-259-1907/Gantry will not power on. Siemens ticket# 400112602478 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC171258 | CT Scanner | 64559 | 399199 | Siemens Medical Solutions USA | SomatomDefinition AS | BOI | BOI
1906426 | 05/14/2024 08:08 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B162030 | Kathryn Gullo / 334.494.4072 | X-Ray table's Lead shield is unravelingEquipment needs tag | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4414 | Angiographic/Interventional Sy | 109463 | 399199 | Siemens Medical Solutions USA | 10848281 | SJS | SJS
1906642 | 05/15/2024 07:15 AM | CM | Corrective Maintenance | DU | Device Unavailable | B165360 | Victoria Perez / 813-507-3032 | 813-302-8045. Cassette is not connecting. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS7168 | Digital Portable X-ray System | DF2401900182WK | 439946 | GE Healthcare USA | 5555000-27 | SJS | SJS
1906745 | 05/15/2024 12:53 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B158750 | Gerson Colon jr / 813-210-6778 | Per Jamie Larussa: "Several issues with the TIMS computer." | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH24247 | Dicom Conversion System | 7001137 | 471654 | TIMS Medical Div Foresight Ima | 061000-2000 | SJH | SJH
1906747 | 05/14/2024 01:01 PM | CM | Corrective Maintenance | WP | Awaiting Parts | B158750 | Gerson Colon jr / 813-210-6778 | Nikki Moore stated: "Ceiling monitors keep moving around." | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS2556 | CT Scanner 128 | 95021 | 399199 | Siemens Medical Solutions USA | 8098027 | SJS | SJS
1906831 | 05/16/2024 03:55 AM | CM | Corrective Maintenance | AF | Awaiting FSR | B090105 | Justin Olson / 727-825-1040 | ASAP - Not turning on after CT machine was reset. 813-302-8851 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS2556 | CT Scanner 128 | 95021 | 399199 | Siemens Medical Solutions USA | 8098027 | SJS | SJS
1911769 | 05/20/2024 10:54 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | dey | dey | Cannot do spot compression view. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH4072 | 2D Digital Mammo | 81005110811 | 106487 | Hologic Inc | SDM-00001-2D | SJH | SJH
1911770 | 06/01/2024 10:55 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Janet Pindral | 813.872.3949 | -Service Agreement PM - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH4072 | 2D Digital Mammo | 81005110811 | 106487 | Hologic Inc | SDM-00001-2D | SJH | SJH
1911815 | 05/20/2024 01:55 PM | CM | Corrective Maintenance | IP | In Progress | B132784 | Deyanira Rizzo / 813-356-7112 | Error code ZTA Call Services completely down. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH4067 | 2D Digital Mammo | 81008110911 | 106487 | Hologic Inc | SDM-00001-2D | SJH | SJH
1912278 | 05/22/2024 04:13 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B110516 | Abby Hastings / 813-586-8511 | 813-586-8511 / Top part has come off, can''t get it back on. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC176686 | 3D Digital Mammo | 3DM160100543 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1913007 | 05/28/2024 07:48 AM | CM | Corrective Maintenance | IP | In Progress | B119731 | Dannielle Alfonso / | 813-259-1907. CT Machine is down. Reference ticket from Siemens 40011262701. | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | BC171258 | CT Scanner | 64559 | 399199 | Siemens Medical Solutions USA | SomatomDefinition AS | BOI | BOI
1913179 | 05/28/2024 01:45 PM | CM | Corrective Maintenance | IP | In Progress | B105822 | Dianne Gee / 813-443-7163 | Will not go down on column, unable to activate drive. Location: Radiology by nuke med dept. Contact: 8133028045 TM called to get update and advised they need device-8133028045 TM called and advised device is working and WO can be closed. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS6405 | Digital Portable X-ray System | 1034415WK8 | 439946 | GE Healthcare USA | 5555000-5 | SJS | SJS
1913191 | 05/28/2024 02:23 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B110516 | Abby Hastings / 813-586-8511 | 813-493-1176. Cad marking are showing up on all images. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC176686 | 3D Digital Mammo | 3DM160100543 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1913192 | 05/28/2024 02:24 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B110516 | Abby Hastings / 813-586-8511 | 813-493-1176. Cad marking on all images. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC146753 | 3D Digital Mammo | 3DM160100701 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1913483 | 06/01/2024 11:38 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | n | n | Uninterruptible Power Systems - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC110005 | Uninterruptible Power Systems | EG161CBC02 | 450901 | Eaton Corp | 9390 | BOI | BOI
1913484 | 06/01/2024 11:38 AM | PM | Planned Maintenance | DO | Device Operational | n | n | -Pm Device Using Mfg Procedure - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC110006 | Laser Imager | 09927014 | 441054 | Konica Minolta Medical Imaging | 873 | BOI | BOI
1913485 | 06/01/2024 11:38 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | n | n | Uninterruptible Power Systems - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC110007 | Uninterruptible Power Systems | EG161CBC07 | 450901 | Eaton Corp | 9390 | BOI | BOI
1913487 | 06/01/2024 11:38 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | n | n | -Service Agreement PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC110016 | General Digital Rad Room | 22714 | 399199 | Siemens Medical Solutions USA | 10281013 | BOI | BOI
1913490 | 06/01/2024 11:38 AM | PM | Planned Maintenance | AC | Acknowledged | n | n | Densitometer - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC119251 | Densitometer B&W Film QC | 112745 | 102741 | X-Rite Inc | 331 | BOI | BOI
1913491 | 06/01/2024 11:38 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Meredith Alaynick | TBD | -Service Agreement PM - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC120211 | Bone Densitometer | 301988M | 106487 | Hologic Inc | Horizon-W | BOI | BOI
1913499 | 06/01/2024 11:38 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Synda McGraph | TBD | -Service Agreement PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | BC146753 | 3D Digital Mammo | 3DM160100701 | 106487 | Hologic Inc | 3DM-SYS-STD | BOI | BOI
1913528 | 06/01/2024 11:38 AM | PM | Planned Maintenance | AC | Acknowledged | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186802 | 3MP Color Med Display | 2590235115 | 392478 | Barco Inc | MDNC-3421 | BOI | BOI
1913530 | 06/01/2024 11:38 AM | PM | Planned Maintenance | AC | Acknowledged | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186803 | 3MP Color Med Display | 2590235112 | 392478 | Barco Inc | MDNC-3421 | BOI | BOI
1913531 | 06/01/2024 11:38 AM | PM | Planned Maintenance | AC | Acknowledged | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186804 | 3MP Color Med Display | 2590134129 | 392478 | Barco Inc | MDNC-3421 | BOI | BOI
1913532 | 06/01/2024 11:38 AM | PM | Planned Maintenance | AC | Acknowledged | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186806 | 3MP Color Med Display | 2590134130 | 392478 | Barco Inc | MDNC-3421 | BOI | BOI
1913533 | 06/01/2024 11:38 AM | PM | Planned Maintenance | AC | Acknowledged | n | n | Diagnostic Medical Display Annual PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186807 | 12MP diagnostic display | 2590232628 | 392478 | Barco Inc | MDMC-12133 | BOI | BOI
1913534 | 06/01/2024 11:38 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | n | n | -Service Agreement PM and Required Annual Password Change - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186808 | Bone Densitometer | 301016M | 106487 | Hologic Inc | Horizon-W | BOI | BOI
1913580 | 05/30/2024 11:55 AM | CM | Corrective Maintenance | IP | In Progress | B105822 | Dianne Gee / 813-443-7163 | 813-302-8045/Images not sending to PACS. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS5900 | Dicom conversion system | 8002143 | 471654 | TIMS Medical Div Foresight Ima | TIMS 2000 SP | SJS | SJS
1914545 | 06/01/2024 07:17 PM | PM | Planned Maintenance | DU | Device Unavailable | Vikki ONeal | 813-302-8030 | -Service Agreement PM and Required Annual Password Change - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS1528 | Nuclear Medicine Work Station | CZC4123Q3K | 439946 | GE Healthcare USA | Xeleris SPECT | SJS | SJS
1914638 | 06/01/2024 07:18 PM | PM | Planned Maintenance | AC | Acknowledged | Val Relente | 813-302-8030 | --- OR Safety Inspection --- - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS2375 | CFD Mobile C-Arm | FBHXXX00114 | 453503 | GE Healthcare Surgery | OEC Elite CFD | SJS | SJS
1914639 | 06/01/2024 07:18 PM | PM | Planned Maintenance | AC | Acknowledged | Val Relente | 813-302-8030 | --- OR Safety Inspection --- - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS2376 | CFD Mobile C-Arm | FBHXXX00237 | 453503 | GE Healthcare Surgery | OEC Elite CFD | SJS | SJS
1914653 | 06/01/2024 07:18 PM | PM | Planned Maintenance | AC | Acknowledged | Vikki ONeal | 813-302-8030 | -Pm Performed By Outside Service - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJS2669 | Avanto Biopsy Breast Array | 051010 | 308478 | Invivo Corp | 104461 | SJS | SJS
1914664 | 06/01/2024 07:18 PM | PM | Planned Maintenance | AC | Acknowledged | Val Relente | 813-302-8030 | --- OR Safety Inspection --- - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS3389 | CFD Mobile C-Arm | FBHXTX00086 | 453503 | GE Healthcare Surgery | OEC Elite CFD | SJS | SJS
1914668 | 06/01/2024 07:18 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Vikki ONeal | 813-302-8030 | -Service Agreement PM and Required Annual Password Change - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS3465 | Nuclear Medicine Work Station | CZC64386XX | 439946 | GE Healthcare USA | Xeleris SPECT | SJS | SJS
1914699 | 06/01/2024 07:18 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Brian Frazier | 813-302-8340 | -Service Agreement PM - Next Scheduled Date: 09/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4415 | Angiographic/Interventional Sy | 121056 | 399199 | Siemens Medical Solutions USA | 10848282 | SJS | SJS
1914717 | 06/01/2024 07:18 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Vikki ONeal | 813-302-8030 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJS4994 | Nuclear Medicine Camera System | CZC62878VC | 439946 | GE Healthcare USA | 5370540-27 | SJS | SJS
1914721 | 06/01/2024 07:18 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Brian Frazier | 813-302-8340 | -Service Agreement PM - Next Scheduled Date: 09/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS5035 | Angiographic/Interventional Sy | 121057 | 399199 | Siemens Medical Solutions USA | 10848282 | SJS | SJS
1914722 | 06/01/2024 07:18 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Val Relente | 813-302-8030 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJS5041 | Digital Portable X-ray System | 1034668WK2 | 439946 | GE Healthcare USA | 5555000-5 | SJS | SJS
1914739 | 06/01/2024 07:18 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Val Relente | 813-302-8030 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJS5488 | Digital Portable X-ray System | 1034329WK1 | 439946 | GE Healthcare USA | 5555000-5 | SJS | SJS
1914769 | 06/01/2024 07:18 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Vikki ONeal | 813-302-8030 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJS6362 | Breast Biopsy System | 82007193261RM | 106487 | Hologic Inc | RM-STLC-00004 | SJS | SJS
1914770 | 06/01/2024 07:18 PM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Val Relente | 813-302-8030 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJS6405 | Digital Portable X-ray System | 1034415WK8 | 439946 | GE Healthcare USA | 5555000-5 | SJS | SJS
1914771 | 06/01/2024 07:18 PM | PM | Planned Maintenance | DU | Device Unavailable | Val Relente | 813-302-8030 | -Vendor PM - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS6419 | Breast Biopsy System | 31503133064 | 106487 | Hologic Inc | MULTICARE | SJS | SJS
1914838 | 06/01/2024 07:19 PM | PM | Planned Maintenance | AC | Acknowledged | Vikki ONeal | 813-302-8030 | -Pm Performed By Outside Service - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJS7182 | Closed MRI Spine Coil | 6754 IVK | 399199 | Siemens Medical Solutions USA | 07579340 | SJS | SJS
1917629 | 06/01/2024 06:49 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Janet Pindral | 813.872.3949 | -Service Agreement PM - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH26487 | Breast Biopsy System | 82010151269RM | 106487 | Hologic Inc | STLC-00004 | SJH | SJH
1917824 | 06/01/2024 06:50 AM | PM | Planned Maintenance | AC | Acknowledged | Janet Pindral | 813.872.3949 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH4070 | Digital Biopsy System | 20536 | 456588 | Faxitron Bioptics Inc | Core Vision | SJH | SJH
1917825 | 06/01/2024 06:50 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Janet Pindral | 813.872.3949 | -Service Agreement PM - Next Scheduled Date: 12/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH4073 | 2D Digital Mammo | 81005110809 | 106487 | Hologic Inc | SDM-00001-2D | SJH | SJH
1917826 | 06/01/2024 06:50 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Janet Pindral | 813.872.3949 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH4075 | Digital Portable X-ray System | 410001B2A004 | 106973 | Shimadzu Medical Systems, Inc | 503-70000-01 | SJH | SJH
1917827 | 06/01/2024 06:50 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Janet Pindral | 813.872.3949 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH4077 | Digital Portable X-ray System | 410001B23008 | 106973 | Shimadzu Medical Systems, Inc | 503-70000-01 | SJH | SJH
1917935 | 06/01/2024 06:51 AM | PM | Planned Maintenance | AC | Acknowledged | Lorraine Sargent | 813.870.4597 | -Pm Device Using Mfg Procedure - Next Scheduled Date: 08/01/2024 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH8157 | CT Scanner | 8DA15X2004 | 101894 | Toshiba America Medical System | TSX-301C-6 | SJH | SJH
1917936 | 06/01/2024 06:51 AM | PM | Planned Maintenance | IW | Awaiting Onsite Vendor Maint | Michele Cimino | 813.870.4726 | -Service Agreement PM - Next Scheduled Date: 06/01/2025 | Gerson Colon Jr | GECOL | EMPLOYEE | 2 | PM Level 2 | SJH8159 | Digital Biopsy System | 30373 | 106487 | Hologic Inc | BIOVISION 10X15 | SJH | SJH
1918909 | 06/03/2024 02:56 PM | CM | Corrective Maintenance | WP | Awaiting Parts | bsh56849 | Beatrice Hawthorne / 813-870-4605 | Work Order to connect this unit to the network. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH11948 | Mobile C-Arm | E2-5316 | 453503 | GE Healthcare Surgery | 9900 ELITE | SJH | SJH
1918927 | 06/03/2024 03:27 PM | CM | Corrective Maintenance | IP | In Progress | B117911 | Steven Kelly / | 813.302.8045 / Portable x-ray detector is not connecting to the portable. There is also a bar going across the images. / First Floor Imaging, | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | SJS6406 | Digital Portable X-ray System | DF2402000764WK | 439946 | GE Healthcare USA | 5555000-27 | SJS | SJS
1918966 | 06/03/2024 08:30 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | b141672 | Joseph Grande / | SJS. MRI. system is down. | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | SJS2555 | MRI 1.5 Tesla | 31418 | 399199 | Siemens Medical Solutions USA | 10018165 | SJS | SJS
1919185 | 06/04/2024 12:34 PM | CM | Corrective Maintenance | IP | In Progress | B115690 | David Klaue / | 813-302-8743/Not connecting to network. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS6416 | Core Integrated Ultrasound | 103001045001 | 453551 | Philips Healthcare | CORE01 | SJS | SJS
1920769 | 06/07/2024 02:28 PM | CM | Corrective Maintenance | IP | In Progress | B117911 | Steven Kelly / | Electrical cord is frayed - OR 2nd Fl - 813-302-8045 | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | SJS2600 | Mobile C-Arm | E2-4423 | 453503 | GE Healthcare Surgery | 9900 ELITE | SJS | SJS
1920805 | 06/07/2024 05:54 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B150179 | Amanda Evett / 813-302-8045 | 813-302-8045 Device stops randomly.Loc: In 30s hallway in the ER | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | SJS7168 | Digital Portable X-ray System | DF2401900182WK | 439946 | GE Healthcare USA | 5555000-27 | SJS | SJS
1921119 | 06/10/2024 04:53 PM | IN | Initial Inspection | AD | Awaiting Delivery of Equipment | b126836 | Jennah Judd / 727-315-5339 | ETA 7/22/2024 ACUSON ORIGIN HYBRID OR 9 510123223 SIEMENS QUOTE ATTACHED | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | NEW-EQUIP | New Equipment ADD | NSNNEW-EQUIPSJH | 000 | Misc-Numbers | New Equipment ADD | SJH | SJH
1922348 | 06/12/2024 07:50 AM | CM | Corrective Maintenance | IP | In Progress | B139330 | Alejandro Babilonia / | 813 302 8045 portable xray - not charging | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | SJS5488 | Digital Portable X-ray System | 1034329WK1 | 439946 | GE Healthcare USA | 5555000-5 | SJS | SJS
1922353 | 06/12/2024 08:18 AM | CM | Corrective Maintenance | IP | In Progress | B153556 | Stefano Piombino / | Teresa Farrell - 8135867520 / Xray modality at the BUC in Riverview that is not sending exams to PACS as of 5pm last night | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | BC176667 | General Digital Rad Room | AFF346-1018 | 171242 | Summit Industries LLC | L501-01 | BMG | BMG
1922381 | 06/12/2024 09:15 AM | CM | Corrective Maintenance | IP | In Progress | B144971 | Teresa Farrell / 813-586-8686 | 813-586-7520/Machine is only sending one image per study. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC176667 | General Digital Rad Room | AFF346-1018 | 171242 | Summit Industries LLC | L501-01 | BMG | BMG
1922565 | 06/13/2024 07:56 AM | CM | Corrective Maintenance | IP | In Progress | B110516 | Abby Hastings / 813-586-8511 | QA app failed it''s check. BLBHH1RD1027DR, apparently the app is for CES. | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | NOTAG | Need Tag Number | NSNNOTAG | 000 | Misc-Numbers | No Tag | BMG | BMG
1922599 | 06/13/2024 09:54 AM | CM | Corrective Maintenance | IP | In Progress | B105822 | Dianne Gee / 813-443-7163 | 813-302-8045. Light to send to patient is out. | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | SJS6405 | Digital Portable X-ray System | 1034415WK8 | 439946 | GE Healthcare USA | 5555000-5 | SJS | SJS
1922838 | 06/14/2024 08:01 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B068323 | Janette Sanford /813.302.8928 | Table not moving forward/backwardSiemans ticket: 400112659278 | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | BC110002 | CT Scanner 128 | 66611 | 399199 | Siemens Medical Solutions USA | 8098027 | BOI | BOI
1923089 | 06/16/2024 10:00 PM | EN | Easy Net | OP | Open | exd53563 | 8133567113 | Plug not working or charging properly. it sparks and the whole portable turns Black with no access to images. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH4077 | Digital Portable X-ray System | 410001B23008 | 106973 | Shimadzu Medical Systems, Inc | 503-70000-01 | SJH | SJH
1923295 | 06/17/2024 04:17 PM | EN | Easy Net | IW | Awaiting Onsite Vendor Maint | B132173 | 813-302-8016 | 813-302-8016 // SJS Breast Center // The brevara remote wont fire the needle. We are unable to do stereo breast biopsies until that time.6/18/24 @ 11:02AM- TM Rebecca Stuart called to request STAT, said they have procedures today. 6/19/24 - Rebecca calling back no acknowledgment for STAT ticket that was placed 2 days ago. | Gerson Colon Jr | GECOL | EMPLOYEE | 8 | 1 Emergency/STAT | SJS7173 | Breast Biopsy System | BVA01201 | 106487 | Hologic Inc | BREV100 | SJS | SJS
1923510 | 06/18/2024 03:04 PM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B116025 | Michele Maningas / 813-302-8743 | 813-302-8743 / Lead shield on a table needs to be sewn back onto the railing because it has been partially torn off. / 2nd Floor, Room 2263, IR Lab 3 / | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJS4414 | Angiographic/Interventional Sy | 109463 | 399199 | Siemens Medical Solutions USA | 10848281 | SJS | SJS
1923628 | 06/19/2024 12:03 PM | CM | Corrective Maintenance | AC | Acknowledged | B088245 | Melissa Monaghan / 813-870-4605 | Error opening dicom assoc - BOI South DEXA AT - 813-302-8910 | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186808 | Bone Densitometer | 301016M | 106487 | Hologic Inc | Horizon-W | BOI | BOI
1923759 | 06/19/2024 11:00 AM | CM | Corrective Maintenance | IW | Awaiting Onsite Vendor Maint | B158750 | Gerson Colon jr / 813-210-6778 | Becca called me and stated that the new Trident HD unit is not sending Images to PACS. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | SJH22326 | Digital Biopsy System | TRI2101509 | 106487 | Hologic Inc | Trident HD | SJH | SJH
1923831 | 06/20/2024 12:03 PM | CM | Corrective Maintenance | WP | Awaiting Parts | b158750 | Gerson Colon jr. / | Help ordering accessories. | Gerson Colon Jr | GECOL | EMPLOYEE | 9 | 2 Routine | BC186779 | MRI 1.5 Tesla | 31351 | 399199 | Siemens Medical Solutions USA | 10018165 | BOI | BOI